## Summary Data
Summary Data
Description of Munition Item Tested | M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
DODIC | A068
Test Location | Aberdeen Proving Ground
Date Background Run Conducted | 8/27/2001 or 9/6/2001
Number of Test Runs Conducted | 2
Test Run Number | 1 | 2 | Background
Date Test Run Conducted | 8/27/2001 or 9/6/2001 | 8/27/2001 or 9/6/2001 | 8/27/2001 or 9/6/2001
No. of Items Tested | 20 | 20 | 0
Net Explosive Weight for One Item (lb) | 0.00384 | 0.00384 | NA
Net Explosive Weight for All Items (lb) | 0.0768 | 0.0768 | NA
Initial Plume Volume (m3) | 14.716 | 14.716 | 14.716
Note on B7: Kaobrian: FTIR and CEM-measured compounds were sampled on both dates. Metals were sampled on 9/6/2001. All other analytes sections were measured on 8/27/2001.
Note on B8: Kaobrian: Note that the test report states that a second day of testing was required to obtain data because some of the samplers failed during the first day. CEM and FTIR-measured compounds were sampled during 4 test runs.
Note on B11: Kaobrian: FTIR and CEM-measured compounds were sampled on both dates. Metals were sampled on 9/6/2001. All other analytes sections were measured on 8/27/2001.
Note on C11: Kaobrian: FTIR and CEM-measured compounds were sampled on both dates. Metals were sampled on 9/6/2001. All other analytes sections were measured on 8/27/2001.
Note on D11: Kaobrian: FTIR and CEM-measured compounds were sampled on both dates. Metals were sampled on 9/6/2001. All other analytes sections were measured on 8/27/2001.

## Final EF SUM
Summary of CO2, Criteria Pollutant, TSP, Toxic Chemical, and Hazardous Air Pollutant Emission Factors
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
CAS # | Pollutant (Official [AP-42] Notation) | Pollutant (Test Notation) | Emission Factor |  | HAP | Toxic | Detection Limit Based Concentration (mg/m3)
|  |  | lb/item | lb/lb NEW
Carbon Dioxide, Criteria Pollutants, and Total Suspended Particulates
124-38-9 | Carbon dioxide | CO2 | 0.00097761 | 0.2545 | No | No | --
630-08-0 | Carbon monoxide | CO | 0.0016183 | 0.42131 | No | No | --
7439-92-1 | Lead | Lead | 2.7632e-06 | 0.00071959 | Yes | Yes | --
74-82-8 | Methane | Methane | 6.2094e-06 | 0.001617 | No | No | --
-- | Oxides of nitrogen | NOx | 1.7461e-05 | 0.0045454 | No | No | --
-- | PM-2.5 | PM2.5 | 5.1225e-05 | 0.01334 | No | No | --
-- | PM-10 | PM10 | 6.7046e-05 | 0.01746 | No | No | --
7446-09-5 | Sulfur dioxide | SO2 | ND | ND | No | No | 0.53248
12789-66-1 | Total suspended particulate | TSP | 7.0157e-05 | 0.01827 | No | No | --
Toxic Chemicals and Hazardous Air Pollutants
83-32-9 | Acenaphthene | Acenaphthene | 1.7354e-10 | 4.5193e-08 | Yes | No | --
208-96-8 | Acenaphthylene | Acenaphthylene | 1.2012e-09 | 3.1281e-07 | Yes | No | --
75-07-0 | Acetaldehyde | Acetaldehyde | 1.231e-07 | 3.2058e-05 | Yes | Yes | --
75-05-8 | Acetonitrile | Acetonitrile | 2.3496e-07 | 6.1187e-05 | Yes | Yes | --
98-86-2 | Acetophenone | Aectophenone | 3.2613e-08 | 8.493e-06 | Yes | Yes | --
107-02-8 | Acrolein | Acrolein | ND | ND | Yes | Yes | 0.01334
107-13-1 | Acrylonitrile | Acrylonitrile | 2.0116e-08 | 5.2384e-06 | Yes | Yes | --
107-05-1 | Allyl chloride | 3-Chloropropene | ND | ND | Yes | Yes | 0.00089142
7429-90-5 | Aluminum | Aluminum | 1.1024e-07 | 2.8708e-05 | No | Yes | --
7664-41-7 | Ammonia | NH3 | 3.7829e-05 | 0.0098479 | Yes | No | --
120-12-7 | Anthracene | Anthracene | 9.1271e-11 | 2.3769e-08 | Yes | Yes | --
7440-36-0 | Antimony | Antimony | 1.325e-06 | 0.00034504 | Yes | Yes | --
7440-38-2 | Arsenic | Arsenic | ND | ND | Yes | Yes | 0.0090448
7440-39-3 | Barium | Barium | 4.6637e-07 | 0.00012145 | No | Yes | --
71-43-2 | Benzene | Benzene | 1.9242e-07 | 5.011e-05 | Yes | Yes | --
56-55-3 | Benzo[a]anthracene | Benzo(a)anthracene | ND | ND | Yes | Yes | 8.1333e-05
205-99-2 | Benzo[b]fluoranthene | Benzo(b)fluoranthene | 1.2108e-10 | 3.1531e-08 | Yes | Yes | --
207-08-9 | Benzo[k]fluoranthene | Benzo(k)fluoranthene | ND | ND | Yes | Yes | 8.1333e-05
191-24-2 | Benzo[g,h,i]perylene | Benzo(g,h,i)perylene | ND | ND | Yes | Yes | 0.016267
50-32-8 | Benzo[a]pyrene | Benzo(a)pyrene | 1.196e-10 | 3.1145e-08 | Yes | Yes | --
192-97-2 | Benzo[e]pyrene | Benzo(e)pyrene | 1.2256e-10 | 3.1917e-08 | Yes | No | --
100-44-7 | Benzyl chloride | Benzylchloride | ND | ND | Yes | Yes | 0.00026416
7440-41-7 | Beryllium | Beryllium | ND | ND | Yes | Yes | 0.0045224
75-25-2 | Bromoform | Bromoform | ND | ND | Yes | Yes | 0.00052568
74-83-9 | Bromomethane | Bromomethane | ND | ND | Yes | Yes | 0.00035546
101-55-3 | 4-Bromophenylphenylether | 4-bromophenyl-phenylether | ND | ND | Yes | Yes | 0.016267
106-99-0 | 1,3-Butadiene | 1,3-Butadiene | ND | ND | Yes | Yes | 0.045003
75-65-0 | t-Butyl alcohol | tert-Butyl Alcohol | ND | ND | Yes | Yes | 0.00040084
85-68-7 | Butylbenzylphthalate | Butylbenzylphthalate | ND | ND | Yes | No | 0.016267
123-72-8 | Butyraldehyde | Butyraldehyde | ND | ND | No | Yes | 0.020524
7440-43-9 | Cadmium | Cadmium | ND | ND | Yes | Yes | 0.0045224
75-15-0 | Carbon disulfide | Carbon Disulfide | ND | ND | Yes | Yes | 0.00031616
56-23-5 | Carbon tetrachloride | Carbontetrachloride | ND | ND | Yes | Yes | 0.00032032
106-47-8 | 4-Chloroaniline | 4-chloroaniline | ND | ND | No | Yes | 0.016267
108-90-7 | Chlorobenzene | Chlorobenzene | 9.6582e-10 | 2.5152e-07 | Yes | Yes | --
75-45-6 | Chlorodifluoromethane | Chlorodifluoromethane | ND | ND | No | Yes | 0.0002518
75-00-3 | Chloroethane | Chloroethane | ND | ND | Yes | Yes | 0.00061733
111-91-1 | bis(2-Chloroethoxy)methane | Bis(2-chloroethoxy)methane | ND | ND | No | Yes | 0.016267
111-44-4 | bis(2-Chloroethyl)ether | Bis(2-chloroethyl)ether | ND | ND | Yes | Yes | 0.016267
67-66-3 | Chloroform | Chloroform | ND | ND | Yes | Yes | 0.00039603
74-87-3 | Chloromethane | Chloromethane | 2.2995e-09 | 5.9884e-07 | Yes | Yes | --
91-58-7 | 2-Chloronaphthalene | 2-chloronaphthalene | ND | ND | Yes | No | 0.016267
95-57-8 | 2-Chlorophenol | 2-chlorophenol | ND | ND | No | Yes | 0.016267
7005-72-3 | 4-Chlorophenyl-phenylether | 4-chlorophenyl-phenylether | ND | ND | Yes | No | 0.016267
7440-47-3 | Chromium | Chromium | ND | ND | Yes | Yes | 0.0022612
18540-29-9 | Hexavalent chromium | Hexavalent Chromium | ND | ND | Yes | Yes | 0.00060545
218-01-9 | Chrysene | Chrysene | ND | ND | Yes | Yes | 8.1333e-05
7440-48-4 | Cobalt | Cobalt | ND | ND | Yes | Yes | 0.0045224
7440-50-8 | Copper | Copper | 2.0743e-05 | 0.0054017 | No | Yes | --
106-44-5 | p-Cresol | 4-methylphenol | ND | ND | Yes | Yes | 0.016267
4170-30-3 | Crotonaldehyde | Crotonaldehyde | ND | ND | No | Yes | 0.020524
98-82-8 | Cumene | Cumene | ND | ND | Yes | Yes | 0.0002
57-12-5 | Particulate cyanide | Particulate Cyanide | 1.2383e-06 | 0.00032247 | Yes | Yes | --
53-70-3 | Dibenz[a,h]anthracene | Dibenz(a,h)anthracene | ND | ND | Yes | Yes | 1.6267e-05
132-64-9 | Dibenzofuran | Dibenzofuran | ND | ND | Yes | Yes | 0.016267
106-93-4 | 1,2-Dibromoethane | 1,2-Dibromoethane | ND | ND | Yes | Yes | 0.00039104
84-74-2 | Dibutyl phthalate | Di-n-butylphthalate | ND | ND | Yes | Yes | 0.016267
95-50-1 | 1,2-Dichlorobenzene | 1,2-Dichlorobenzene | ND | ND | No | Yes | 0.00048922
541-73-1 | 1,3-Dichlorobenzene | 1,3-Dichlorobenzene | ND | ND | No | Yes | 0.00042806
106-46-7 | 1,4-Dichlorobenzene | 1,4-Dichlorobenzene | ND | ND | Yes | Yes | 0.00042806
75-27-4 | Dichlorobromomethane | Bromodichloromethane | ND | ND | No | Yes | 0.00040892
75-71-8 | Dichlorodifluoromethane | Dichlorodifluoromethane | 0 | 0 | No | Yes | --
75-34-3 | 1,1-Dichloroethane | 1,1-Dichloroethane | ND | ND | Yes | Yes | 0.00037066
107-06-2 | 1,2-Dichloroethane | 1,2-Dichloroethane | 3.7574e-09 | 9.785e-07 | Yes | Yes | --
75-43-4 | Dichlorofluoromethane | Dichlorofluoromethane | ND | ND | No | Yes | 0.00021407
120-83-2 | 2,4-Dichlorophenol | 2,4-dichlorophenol | ND | ND | No | Yes | 0.016267
78-87-5 | 1,2-Dichloropropane | 1,2-Dichloropropane | ND | ND | Yes | Yes | 0.00032906
10061-02-6 | trans-1,3-Dichloropropene | trans-1,3-Dichloropropene | ND | ND | No | Yes | 0.00023082
76-14-2 | Dichlorotetrafluoroethane | Freon 114 | ND | ND | No | Yes | 0.00049775
105-67-9 | 2,4-Dimethylphenol | 2,4-dimethylphenol | ND | ND | No | Yes | 0.016267
131-11-3 | Dimethyl phthalate | Dimethylphthalate | ND | ND | Yes | Yes | 0.016267
99-65-0 | 1,3-Dinitrobenzene | 1,3-Dinitrobenzene | ND | ND | No | Yes | 0.00067909
534-52-1 | 4,6-Dinitro-2-methylphenol | 4,6-dinitro-2-methylphenol | ND | ND | Yes | Yes | 0.016267
51-28-5 | 2,4-Dinitrophenol | 2,4-dinitrophenol | ND | ND | Yes | Yes | 0.016267
121-14-2 | 2,4-Dinitrotoluene | 2,4-Dinitrotoluene | ND | ND | Yes | Yes | 0.00067909
606-20-2 | 2,6-Dinitrotoluene | 2,6-Dinitrotoluene | ND | ND | No | Yes | 0.00067909
123-91-1 | 1,4-Dioxane | 1,4-Dioxane | ND | ND | Yes | Yes | 0.00032988
-- | Total dioxin/furan compounds | Total Dixoin | 1.6691e-13 | 4.3466e-11 | Yes | Yes | --
140-88-5 | Ethyl acrylate | Ethyl Acrylate | ND | ND | Yes | Yes | 0.0001666
100-41-4 | Ethylbenzene | Ethylbenzene | 1.6118e-09 | 4.1975e-07 | Yes | Yes | --
74-85-1 | Ethylene | Ethylene | 3.7792e-07 | 9.8418e-05 | No | Yes | --
117-81-7 | bis(2-Ethylhexyl)phthalate | bis(2-Ethylhexyl)phthalate | 2.9849e-08 | 7.7732e-06 | Yes | Yes | --
206-44-0 | Fluoranthene | Fluoranthene | ND | ND | Yes | Yes | 0.016267
86-73-7 | Fluorene | Fluorene | 4.4095e-10 | 1.1483e-07 | Yes | No | --
50-00-0 | Formaldehyde | Formaldehyde | 1.5505e-07 | 4.0376e-05 | Yes | Yes | --
76-13-1 | Freon 113 | Freon 113 | ND | ND | No | Yes | 0.00054696
35822-46-9 | 1,2,3,4,6,7,8-Heptachlorodibenzo-p-dioxin | 1,2,3,4,6,7,8-HpCDD | 4.0353e-14 | 1.0508e-11 | Yes | Yes | --
67562-39-4 | 1,2,3,4,6,7,8-Heptachlorodibenzofuran | 1,2,3,4,6,7,8-HpCDF | 2.9702e-14 | 7.7348e-12 | Yes | Yes | --
55673-89-7 | 1,2,3,4,7,8,9-Heptachlorodibenzofuran | 1234789-HPCDF | 0 | 0 | Yes | Yes | --
118-74-1 | Hexachlorobenzene | Hexachlorobenzene | ND | ND | Yes | Yes | 0.016267
87-68-3 | Hexachlorobutadiene | Hexachlorobutadiene | ND | ND | Yes | Yes | 0.00076003
77-47-4 | Hexacholorocyclopentadiene | Hexachlorocyclopentadiene | ND | ND | Yes | Yes | 0.016267
39227-28-6 | 1,2,3,4,7,8-Hexachlorodibenzo-p-dioxin | 123478-HXCDD | 0 | 0 | Yes | Yes | --
57653-85-7 | 1,2,3,6,7,8-Hexachlorodibenzo-p-dioxin | 123678-HXCDD | 1.0152e-14 | 2.6436e-12 | Yes | Yes | --
19408-74-3 | 1,2,3,7,8,9-Hexachlorodibenzo-p-dioxin | 123789-HXCDD | 6.2003e-17 | 1.6147e-14 | Yes | Yes | --
70648-26-9 | 1,2,3,4,7,8-Hexachlorodibenzofuran | 1,2,3,4,7,8-HxCDF | 1.0018e-14 | 2.6089e-12 | Yes | Yes | --
57117-44-9 | 1,2,3,6,7,8-Hexachlorodibenzofuran | 1,2,3,6,7,8-HxCDF | 2.4305e-15 | 6.3294e-13 | Yes | Yes | --
72918-21-9 | 1,2,3,7,8,9-Hexachlorodibenzofuran | 123789-HXCDF | ND | ND | Yes | Yes | 3.6854e-09
60851-34-5 | 2,3,4,6,7,8-Hexachlorodibenzofuran | 2,3,4,6,7,8-HxCDF | 0 | 0 | Yes | Yes | --
67-72-1 | Hexachloroethane | Hexachlorethane | ND | ND | Yes | Yes | 0.0005909
110-54-3 | Hexane | n-Hexane | ND | ND | Yes | Yes | 0.071702
7647-01-0 | Hydrochloric acid | Hydrogen Chloride | ND | ND | Yes | Yes | 0.17452
74-90-8 | Hydrogen cyanide | Hydrogen Cyanide | 2.9718e-06 | 0.00077389 | Yes | Yes | --
7664-39-3 | Hydrogen fluoride | Hydrogen Fluoride | ND | ND | Yes | Yes | 0.18478
193-39-5 | Indeno[1,2,3-cd]pyrene | Indeno(1,2,3-cd)pyrene | ND | ND | Yes | Yes | 0.016267
78-59-1 | Isophorone | Isophorone | ND | ND | Yes | No | 0.016267
7439-92-1 | Lead | Lead | 2.7632e-06 | 0.00071959 | Yes | Yes | --
7439-96-5 | Manganese | Manganese | ND | ND | Yes | Yes | 0.0045224
96-33-3 | Methyl acrylate | Methyl Acrylate | ND | ND | No | Yes | 0.00025072
74-95-3 | Methylene bromide | Dibromomethane | ND | ND | No | Yes | 0.00050619
75-09-2 | Methylene chloride | Methylene chloride | 0 | 0 | Yes | Yes | --
78-93-3 | Methyl ethyl ketone | 2-Butanone | 3.8245e-09 | 9.9597e-07 | Yes | Yes | --
74-88-4 | Methyl iodide | Methyl Iodide | ND | ND | Yes | Yes | 0.00035428
108-10-1 | Methyl isobutyl ketone | 4-Methyl-2-pentanone | ND | ND | Yes | Yes | 0.0003328
80-62-6 | Methyl methacrylate | Methyl Methacrylate | ND | ND | Yes | Yes | 0.00045815
91-57-6 | 2-Methylnaphthalene | 2-Methylnaphthalene | ND | ND | Yes | No | 0.016267
95-48-7 | 2-Methylphenol | 2-methylphenol | ND | ND | Yes | Yes | 0.016267
1634-04-4 | Methyl tert-butyl ether | Methyl t-Butyl Ether | 6.6837e-10 | 1.7406e-07 | Yes | Yes | --
91-20-3 | Naphthalene | Naphthalene | 1.2712e-08 | 3.3104e-06 | Yes | Yes | --
7440-02-0 | Nickel | Nickel | ND | ND | Yes | Yes | 0.0090448
7697-37-2 | Nitric acid | Nitric Acid | 2.4809e-07 | 6.4606e-05 | No | Yes | --
100-01-6 | 4-Nitroaniline | 4-nitroaniline | ND | ND | No | Yes | 0.016267
98-95-3 | Nitrobenzene | Nitrobenzene | ND | ND | Yes | Yes | 0.00067909
55-63-0 | Nitroglycerin | Nitroglycerine | ND | ND | No | Yes | 0.00069722
88-75-5 | 2-Nitrophenol | 2-nitrophenol | ND | ND | No | Yes | 0.016267
100-02-7 | 4-Nitrophenol | 4-nitrophenol | ND | ND | Yes | Yes | 0.016267
62-75-9 | n-Nitrosodimethylamine | N-nitrosodimethylamine | ND | ND | Yes | Yes | 0.016267
86-30-6 | n-Nitrosodiphenylamine | N-nitrosodiphenylamine(1) | ND | ND | Yes | Yes | 0.016267
621-64-7 | n-Nitroso-di-n-propylamine | N-nitroso-di-n-propylamine | ND | ND | No | Yes | 0.016267
3268-87-9 | 1,2,3,4,6,7,8,9-Octachlorodibenzo-p-dioxin | OCDD | 0 | 0 | Yes | Yes | --
39001-02-0 | 1,2,3,4,6,7,8,9-Octachlorodibenzofuran | OCDF | 0 | 0 | Yes | Yes | --
40321-76-4 | 1,2,3,7,8-Pentachlorodibenzo-p-dioxin | 12378-PECDD | 4.0302e-15 | 1.0495e-12 | Yes | Yes | --
57117-41-6 | 1,2,3,7,8-Pentachlorodibenzofuran | 12378-PECDF | 9.2835e-15 | 2.4176e-12 | Yes | Yes | --
57117-31-4 | 2,3,4,7,8-Pentachlorodibenzofuran | 23478-PECDF | 1.0192e-14 | 2.6541e-12 | Yes | Yes | --
87-86-5 | Pentachlorophenol | Pentachlorophenol | ND | ND | Yes | Yes | 0.016267
85-01-8 | Phenanthrene | Phenanthrene | ND | ND | Yes | Yes | 0.016267
108-95-2 | Phenol | Phenol | ND | ND | Yes | Yes | 0.016267
123-38-6 | Propionaldehyde | Proprionaldehyde | ND | ND | Yes | Yes | 0.020524
115-07-1 | Propylene | Propylene | 5.2681e-08 | 1.3719e-05 | No | Yes | --
129-00-0 | Pyrene | Pyrene | ND | ND | Yes | No | 0.016267
7782-49-2 | Selenium | Selenium | ND | ND | Yes | Yes | 0.0090448
7440-22-4 | Silver | Silver | ND | ND | No | Yes | 0.0022612
100-42-5 | Styrene | Styrene | 7.8944e-09 | 2.0558e-06 | Yes | Yes | --
7664-93-9 | Sulfuric acid | Sulfuric Acid | 3.5128e-07 | 9.148e-05 | No | Yes | --
1746-01-6 | 2,3,7,8-Tetrachlorodibenzo-p-dioxin | 2378-TCDD | 2.9159e-14 | 7.5935e-12 | Yes | Yes | --
51207-31-9 | 2,3,7,8-Tetrachlorodibenzofuran | 2,3,7,8-TCDF | 2.153e-14 | 5.6068e-12 | Yes | Yes | --
630-20-6 | 1,1,1,2-Tetrachloroethane | 1,1,1,2-Tetrachloroethane | ND | ND | No | Yes | 0.00034913
79-34-5 | 1,1,2,2-Tetrachloroethane | 1,1,2,2-Tetrachloroethane | ND | ND | Yes | Yes | 0.00034944
127-18-4 | Tetrachloroethylene | Tertrachloroethene | ND | ND | Yes | Yes | 0.00041391
7440-28-0 | Thallium | Thallium | 1.3748e-08 | 3.5803e-06 | No | Yes | --
108-88-3 | Toluene | Toluene | 1.536e-08 | 3.9999e-06 | Yes | Yes | --
120-82-1 | 1,2,4-Trichlorobenzene | 1,2,4-Trichlorobenzene | ND | ND | Yes | Yes | 0.00128
71-55-6 | 1,1,1-Trichloroethane | 1,1,1-Trichloroethane | ND | ND | Yes | Yes | 0.00044399
79-00-5 | 1,1,2-Trichloroethane | 1,1,2-Trichloroethane | ND | ND | Yes | Yes | 0.00027664
75-69-4 | Trichlorofluoromethane | Trichloroflouromethane | 0 | 0 | No | Yes | --
95-95-4 | 2,4,5-Trichlorophenol | 2,4,5-trichlorophenol | ND | ND | Yes | Yes | 0.016267
88-06-2 | 2,4,6-Trichlorophenol | 2,4,6-trichlorophenol | ND | ND | Yes | Yes | 0.016267
96-18-4 | 1,2,3-Trichloropropane | 1,2,3-Trichloropropane | ND | ND | No | Yes | 0.00049065
95-63-6 | 1,2,4-Trimethylbenzene | 1,2,4-Trimethylbenzene | 0 | 0 | No | Yes | --
540-84-1 | 2,2,4-Trimethylpentane | Isooctane | ND | ND | Yes | No | 0.00019008
7440-62-2 | Vanadium | Vanadium | ND | ND | No | Yes | 0.0045224
108-05-4 | Vinyl acetate | Vinyl Acetate | ND | ND | Yes | Yes | 0.00039395
75-01-4 | Vinyl chloride | Vinyl Chloride | ND | ND | Yes | Yes | 0.00026
75-35-4 | Vinylidene chloride | 1,1-Dichlorethene | ND | ND | Yes | Yes | 0.00020584
106-42-3, 108-38-3 | m-Xylene, p-Xylene | m&p-Xylene | 0 | 0 | Yes | Yes | --
95-47-6 | o-Xylene | o-Xylene | 0 | 0 | Yes | Yes | --
7440-66-6 | Zinc | Zinc | 2.3808e-06 | 0.00062001 | No | Yes | --
Other Pollutants
67-64-1 | Acetone | Acetone | 0 | 0 | No | No | --
74-86-2 | Acetylene | Acetylene | 5.4359e-08 | 1.4156e-05 | No | No | --
35572-78-2 | 2-Amino-4,6-dinitrotoluene | 2-Amino-4,6-Dinitrotoluene | ND | ND | No | No | 0.00067909
19406-51-0 | 4-Amino-2,6-dinitrotoluene | 4-Amino-2,6-Dinitrotoluene | ND | ND | No | No | 0.00067909
100-52-7 | Benzaldehyde | Benzaldehyde | ND | ND | No | No | 0.020524
4748-78-1 | Benzaldehyde, 4-ethyl- | Benzaldehyde, 4-ethyl- | 2.2533e-08 | 5.8679e-06 | No | No | --
100-51-6 | Benzyl alcohol | Benzyl alcohol | ND | ND | No | No | 0.016267
108-86-1 | Bromobenzene | Bromobenzene | ND | ND | No | No | 0.0003919
106-97-8 | Butane | Butane | ND | ND | No | No | 0.053191
106-98-9 | 1-Butene | 1-Butene | ND | ND | No | No | 0.046684
590-18-1 | cis-Butene | cis-butene | ND | ND | No | No | 0.046692
624-64-6 | trans-Butene | trans-Butene | ND | ND | No | No | 0.046692
107-00-6 | 1-Butyne | 1-Butyne | ND | ND | No | No | 0.045003
503-17-3 | 2-Butyne | 2-Butyne | ND | ND | No | No | 0.045003
7440-70-2 | Calcium | Calcium | 2.3022e-07 | 5.9954e-05 | No | No | --
7440-44-0 | Elemental carbon | Elemental Carbon | 9.2963e-07 | 0.00024209 | No | No | --
7440-44-0 | Organic carbon | Organic Carbon | 3.9123e-06 | 0.0010188 | No | No | --
7440-44-0 | Total carbon | Total Carbon | 4.842e-06 | 0.0012609 | No | No | --
39638-32-9 | bis(2-Chloroisopropyl)ether | Bis(2-chloroisopropyl)ether | ND | ND | No | No | 0.016267
59-50-7 | 4-Chloro-3-methylphenol | 4-chloro-3-methylphenol | ND | ND | No | No | 0.016267
124-48-1 | Dibromochloromethane | Dibromochloromethane | ND | ND | No | No | 0.00051969
156-59-2 | cis-1,2-Dichloroethene | cis-1,2-Dichloroethene | ND | ND | No | No | 0.00020164
156-60-5 | trans-1,2-Dichloroethene | trans-1,2,-Dichloroethene | ND | ND | No | No | 0.00048422
10061-01-5 | cis-1,3-Dichloropropylene | cis-1,3-Dichloropropene | ND | ND | No | No | 0.00027698
84-66-2 | Diethylphthalate | Diethylphthalate | ND | ND | No | No | 0.016267
5779-94-2 | 2,5-Dimethylbenzaldehyde | 2,5-Dimethylbenzaldehyde | ND | ND | No | No | 0.020524
117-84-0 | Di-n-octylphthalate | Di-n-octylphthalate | ND | ND | No | No | 0.016267
112-95-8 | Eicosane | Eicosane | 1.3638e-08 | 3.5516e-06 | No | No | --
74-84-0 | Ethane | Ethane | 1.9515e-07 | 5.0821e-05 | No | No | --
64-17-5 | Ethanol | Ethanol | 0 | 0 | No | No | --
141-78-6 | Ethyl acetate | Ethyl Acetate | ND | ND | No | No | 0.0013562
97-63-2 | Ethyl methacrylate | Ethyl Methacrylate | ND | ND | No | No | 0.00033238
622-96-8 | 4-Ethyltoluene | 4-Ethyltoluene | 0 | 0 | No | No | --
142-82-5 | n-Heptane | Heptane | ND | ND | No | No | 0.00020842
66-25-1 | Hexaldehyde | Hexaldehyde | ND | ND | No | No | 0.020524
591-78-6 | 2-Hexanone | 2-Hexanone | ND | ND | No | No | 0.0002912
2691-41-0 | HMX | HMX | ND | ND | No | No | 0.0067909
10035-10-6 | Hydrogen bromide | Hydrogen Bromide | ND | ND | No | No | 0.17452
123-92-2 | Isoamyl acetate | 1-Butanol, 3-methyl-, acetate | 1.6327e-08 | 4.2519e-06 | No | No | --
75-28-5 | Isobutane | Isobutane | ND | ND | No | No | 0.048356
626-19-7 | Isophthalaldehyde | Isophthalaldehyde | 3.5578e-08 | 9.2651e-06 | No | No | --
590-86-3 | Isovaleraldehyde | Isovaleraldehyde | ND | ND | No | No | 0.020524
7439-95-4 | Magnesium | Magnesium | 4.195e-07 | 0.00010925 | No | No | --
98-83-9 | alpha-Methylstyrene | Alpha Methyl Styrene | ND | ND | No | No | 0.00029498
479-45-8 | Methyl-2,4,6-trinitrophenylnitramine | Tetryl | ND | ND | No | No | 0.00067909
88-74-7 | 2-Nitroaniline | 2-nitroaniline | ND | ND | No | No | 0.016267
99-09-2 | 3-Nitroaniline | 3-nitroaniline | ND | ND | No | No | 0.016267
88-72-2 | 2-Nitrotoluene | 2-Nitrotoluene | ND | ND | No | No | 0.00067909
99-08-1 | 3-Nitrotoluene | 3-Nitrotoluene | ND | ND | No | No | 0.00067909
99-99-0 | 4-Nitrotoluene | 4-Nitrotoluene | ND | ND | No | No | 0.00067909
111-65-9 | Octane | Octane | ND | ND | No | No | 0.00019008
109-66-0 | Pentane | n-Pentane | ND | ND | No | No | 0.060029
7664-38-2 | Phosphoric acid | Phosphoric Acid | ND | ND | No | No | 0.17452
74-98-6 | Propane | Propane | ND | ND | No | No | 0.036691
115-11-7 | 1-Propene, 2-methyl | Isobutylene | ND | ND | No | No | 0.046675
74-99-7 | Propyne | Propyne | ND | ND | No | No | 0.032531
121-82-4 | RDX | RDX | ND | ND | No | No | 0.00067909
1334-78-7 | o,m,p-Tolualdehyde | o,m,p-Tolualdehyde | ND | ND | No | No | 0.061571
25323-89-1 | Trichloroethane | Trichloroethane | ND | ND | No | No | 0.00029797
108-67-8 | 1,3,5-Trimethylbenzene | 1,3,5-Trimethylbenzene | 0 | 0 | No | No | --
99-35-4 | 1,3,5-Trinitrobenzene | 1,3,5-Trinitrobenzene | ND | ND | No | No | 0.00067909
118-96-7 | 2,4,6-Trinitrotoluene | 2,4,6-Trinitrotoluene | ND | ND | No | No | 0.00067909
110-62-3 | Valeraldehyde | Valeraldehyde | ND | ND | No | No | 0.020524
NOTE: Blue shading indicates poor precision. These compounds had a relative percent difference greater than 100% within a test.
NOTE : Orange highlights indicate Tentatively Identified Compounds (TICs). TICs will be downgraded in AP-42.
NOTE : Lb/lb NEW factors may be slightly different for the same lb/item value due to rounding.
NOTE : Purple highlights indicate compounds present in the lab blank at levels between 20 and 50 % of test values. These compounds need to be annotated in AP-42 in the following way. These emission factors are assumed to be biased high.
NOTE: Teal highlights indicate the holding time allowance was significantly exceeded prior to analysis. Compounds will be downgraded in AP-42.
AP-42 letter notation:
d Pollutant is listed as a reportable chemical under Section 313 of the Emergency Planning and Community Right-to-Know Act (EPCRA) and as a hazardous air pollutant under Section 112 of the 1990 Clean Air Act Amendments (CAAA).
e Pollutant is listed as a reportable chemical under Section 313 of EPCRA.
f Pollutant is listed as a hazardous air pollutant under Section 112 of the CAAA.
Note on F10: Kaobrian: lead compounds
Note on F18: Kaobrian: POM
Note on F19: Kaobrian: POM
Note on F28: Kaobrian: POM
Note on F29: Kaobrian: "antimony compounds"
Note on F30: Kaobrian: Arsenic Compounds
Note on F33: Kaobrian: POM
Note on F34: Kaobrian: POM
Note on F35: Kaobrian: POM
Note on F36: Kaobrian: POM
Note on F37: Kaobrian: POM
Note on F38: Kaobrian: POM
Note on F40: Kaobrian: Beryllium Compounds
Note on F43: Kaobrian: POM
Note on F46: Kaobrian: POM
Note on F48: Kaobrian: cadmium compounds
Note on F59: Kaobrian: POM
Note on F61: Kaobrian: POM
Note on F62: Kaobrian: Chromium compounds
Note on F64: Kaobrian: POM
Note on F65: Kaobrian: cobalt compounds
Note on F70: Kaobrian: cyanide compounds
Note on F71: Kaobrian: POM
Note on F101: Kaobrian: POM
Note on F120: Kaobrian: cyanide compounds
Note on F125: Kaobrian: manganese compounds
Note on F133: Kaobrian: POM
Note on F137: Kaobrian: Nickel compounds
Note on F145: Kaobrian: POM
Note on F157: Kaobrian: POM
Note on F158: Kaobrian: Selenium Compounds
Note on F162: Kaobrian: POM

## EF SUM with duplicates
Summary of CO2, Criteria Pollutant, TSP, Toxic Chemical, and Hazardous Air Pollutant Emission Factors
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
CAS # | Pollutant (Official [AP-42] Notation) | Pollutant (Test Notation) | Emission Factor |  | HAP | Toxic | Detection Limit Based Concentration (mg/m3)
|  |  | lb/item | lb/lb NEW
Carbon Dioxide, Criteria Pollutants, and Total Suspended Particulates
124-38-9 | Carbon dioxide | CO2 | 0.00097761 | 0.2545 | No | No | --
630-08-0 | Carbon monoxide | CO | 0.0016183 | 0.42131 | No | No | --
7439-92-1 | Lead | Lead | 2.7632e-06 | 0.00071959 | Yes | Yes | --
74-82-8 | Methane | Methane | 6.2094e-06 | 0.001617 | No | No | --
-- | Oxides of nitrogen | NOx | 1.7461e-05 | 0.0045454 | No | No | --
-- | PM-2.5 | PM2.5 | 5.1225e-05 | 0.01334 | No | No | --
-- | PM-10 | PM10 | 6.7046e-05 | 0.01746 | No | No | --
7446-09-5 | Sulfur dioxide | SO2 | ND | ND | No | No | 0.53248
12789-66-1 | Total suspended particulate | TSP | 7.0157e-05 | 0.01827 | No | No | --
Toxic Chemicals and Hazardous Air Pollutants
83-32-9 | Acenaphthene | Acenaphthene | ND | ND | Yes | No | 0.016267
83-32-9 | Acenaphthene | Acenaphthene | 1.7354e-10 | 4.5193e-08 | Yes | No | --
208-96-8 | Acenaphthylene | Acenaphthylene | ND | ND | Yes | No | 0.016267
208-96-8 | Acenaphthylene | Acenaphthylene | 1.2012e-09 | 3.1281e-07 | Yes | No | --
75-07-0 | Acetaldehyde | Acetaldehyde | 1.231e-07 | 3.2058e-05 | Yes | Yes | --
75-05-8 | Acetonitrile | Acetonitrile | 2.3496e-07 | 6.1187e-05 | Yes | Yes | --
98-86-2 | Acetophenone | Aectophenone | 3.2613e-08 | 8.493e-06 | Yes | Yes | --
107-02-8 | Acrolein | Acrolein | ND | ND | Yes | Yes | 0.01334
107-02-8 | Acrolein | Acrolein | 2.6472e-09 | 6.8937e-07 | Yes | Yes | --
107-13-1 | Acrylonitrile | Acrylonitrile | 2.0116e-08 | 5.2384e-06 | Yes | Yes | --
107-05-1 | Allyl chloride | 3-Chloropropene | ND | ND | Yes | Yes | 0.00089142
7429-90-5 | Aluminum | Aluminum | 1.1024e-07 | 2.8708e-05 | No | Yes | --
7664-41-7 | Ammonia | NH3 | 3.7829e-05 | 0.0098479 | Yes | No | --
120-12-7 | Anthracene | Anthracene | ND | ND | Yes | Yes | 0.016267
120-12-7 | Anthracene | Anthracene | 9.1271e-11 | 2.3769e-08 | Yes | Yes | --
7440-36-0 | Antimony | Antimony | 1.325e-06 | 0.00034504 | Yes | Yes | --
7440-38-2 | Arsenic | Arsenic | ND | ND | Yes | Yes | 0.0090448
7440-39-3 | Barium | Barium | 4.6637e-07 | 0.00012145 | No | Yes | --
71-43-2 | Benzene | Benzene | 1.9242e-07 | 5.011e-05 | Yes | Yes | --
56-55-3 | Benzo[a]anthracene | Benzo(a)anthracene | ND | ND | Yes | Yes | 0.016267
56-55-3 | Benzo[a]anthracene | Benzo(a)anthracene | ND | ND | Yes | Yes | 8.1333e-05
205-99-2 | Benzo[b]fluoranthene | Benzo(b)fluoranthene | ND | ND | Yes | Yes | 0.016267
205-99-2 | Benzo[b]fluoranthene | Benzo(b)fluoranthene | 1.2108e-10 | 3.1531e-08 | Yes | Yes | --
207-08-9 | Benzo[k]fluoranthene | Benzo(k)fluoranthene | ND | ND | Yes | Yes | 0.016267
207-08-9 | Benzo[k]fluoranthene | Benzo(k)fluoranthene | ND | ND | Yes | Yes | 8.1333e-05
191-24-2 | Benzo[g,h,i]perylene | Benzo(g,h,i)perylene | ND | ND | Yes | Yes | 0.016267
191-24-2 | Benzo[g,h,i]perylene | Benzo(g,h,i)perylene | 2.0033e-10 | 5.2169e-08 | Yes | Yes | --
50-32-8 | Benzo[a]pyrene | Benzo(a)pyrene | ND | ND | Yes | Yes | 0.016267
50-32-8 | Benzo[a]pyrene | Benzo(a)pyrene | 1.196e-10 | 3.1145e-08 | Yes | Yes | --
192-97-2 | Benzo[e]pyrene | Benzo(e)pyrene | 1.2256e-10 | 3.1917e-08 | Yes | No | --
100-44-7 | Benzyl chloride | Benzylchloride | ND | ND | Yes | Yes | 0.00026416
7440-41-7 | Beryllium | Beryllium | ND | ND | Yes | Yes | 0.0045224
75-25-2 | Bromoform | Bromoform | ND | ND | Yes | Yes | 0.00052568
74-83-9 | Bromomethane | Bromomethane | ND | ND | Yes | Yes | 0.00035546
101-55-3 | 4-Bromophenylphenylether | 4-bromophenyl-phenylether | ND | ND | Yes | Yes | 0.016267
106-99-0 | 1,3-Butadiene | 1,3-Butadiene | 1.0245e-08 | 2.668e-06 | Yes | Yes | --
106-99-0 | 1,3-Butadiene | 1,3-Butadiene | ND | ND | Yes | Yes | 0.045003
75-65-0 | t-Butyl alcohol | tert-Butyl Alcohol | ND | ND | Yes | Yes | 0.00040084
85-68-7 | Butylbenzylphthalate | Butylbenzylphthalate | ND | ND | Yes | No | 0.016267
123-72-8 | Butyraldehyde | Butyraldehyde | ND | ND | No | Yes | 0.020524
7440-43-9 | Cadmium | Cadmium | ND | ND | Yes | Yes | 0.0045224
75-15-0 | Carbon disulfide | Carbon Disulfide | ND | ND | Yes | Yes | 0.00031616
56-23-5 | Carbon tetrachloride | Carbontetrachloride | ND | ND | Yes | Yes | 0.00032032
106-47-8 | 4-Chloroaniline | 4-chloroaniline | ND | ND | No | Yes | 0.016267
108-90-7 | Chlorobenzene | Chlorobenzene | 9.6582e-10 | 2.5152e-07 | Yes | Yes | --
75-45-6 | Chlorodifluoromethane | Chlorodifluoromethane | ND | ND | No | Yes | 0.0002518
75-00-3 | Chloroethane | Chloroethane | ND | ND | Yes | Yes | 0.00061733
111-91-1 | bis(2-Chloroethoxy)methane | Bis(2-chloroethoxy)methane | ND | ND | No | Yes | 0.016267
111-44-4 | bis(2-Chloroethyl)ether | Bis(2-chloroethyl)ether | ND | ND | Yes | Yes | 0.016267
67-66-3 | Chloroform | Chloroform | ND | ND | Yes | Yes | 0.00039603
74-87-3 | Chloromethane | Chloromethane | 2.2995e-09 | 5.9884e-07 | Yes | Yes | --
91-58-7 | 2-Chloronaphthalene | 2-chloronaphthalene | ND | ND | Yes | No | 0.016267
95-57-8 | 2-Chlorophenol | 2-chlorophenol | ND | ND | No | Yes | 0.016267
7005-72-3 | 4-Chlorophenyl-phenylether | 4-chlorophenyl-phenylether | ND | ND | Yes | No | 0.016267
7440-47-3 | Chromium | Chromium | ND | ND | Yes | Yes | 0.0022612
18540-29-9 | Hexavalent chromium | Hexavalent Chromium | ND | ND | Yes | Yes | 0.00060545
218-01-9 | Chrysene | Chrysene | ND | ND | Yes | Yes | 0.016267
218-01-9 | Chrysene | Chrysene | ND | ND | Yes | Yes | 8.1333e-05
7440-48-4 | Cobalt | Cobalt | ND | ND | Yes | Yes | 0.0045224
7440-50-8 | Copper | Copper | 2.0743e-05 | 0.0054017 | No | Yes | --
106-44-5 | p-Cresol | 4-methylphenol | ND | ND | Yes | Yes | 0.016267
4170-30-3 | Crotonaldehyde | Crotonaldehyde | ND | ND | No | Yes | 0.020524
98-82-8 | Cumene | Cumene | ND | ND | Yes | Yes | 0.0002
57-12-5 | Particulate cyanide | Particulate Cyanide | 1.2383e-06 | 0.00032247 | Yes | Yes | --
53-70-3 | Dibenz[a,h]anthracene | Dibenz(a,h)anthracene | ND | ND | Yes | Yes | 0.016267
53-70-3 | Dibenz[a,h]anthracene | Dibenz(a,h)anthracene | ND | ND | Yes | Yes | 1.6267e-05
132-64-9 | Dibenzofuran | Dibenzofuran | ND | ND | Yes | Yes | 0.016267
106-93-4 | 1,2-Dibromoethane | 1,2-Dibromoethane | ND | ND | Yes | Yes | 0.00039104
84-74-2 | Dibutyl phthalate | Di-n-butylphthalate | ND | ND | Yes | Yes | 0.016267
95-50-1 | 1,2-Dichlorobenzene | 1,2-Dichlorobenzene | ND | ND | No | Yes | 0.00048922
95-50-1 | 1,2-Dichlorobenzene | 1,2-dichlorobenzene | ND | ND | No | Yes | 0.016267
541-73-1 | 1,3-Dichlorobenzene | 1,3-Dichlorobenzene | ND | ND | No | Yes | 0.00042806
541-73-1 | 1,3-Dichlorobenzene | 1,3-dichlorobenzene | ND | ND | No | Yes | 0.016267
106-46-7 | 1,4-Dichlorobenzene | 1,4-Dichlorobenzene | ND | ND | Yes | Yes | 0.00042806
106-46-7 | 1,4-Dichlorobenzene | 1,4-dichlorobenzene | ND | ND | Yes | Yes | 0.016267
75-27-4 | Dichlorobromomethane | Bromodichloromethane | ND | ND | No | Yes | 0.00040892
75-71-8 | Dichlorodifluoromethane | Dichlorodifluoromethane | 0 | 0 | No | Yes | --
75-34-3 | 1,1-Dichloroethane | 1,1-Dichloroethane | ND | ND | Yes | Yes | 0.00037066
75-43-4 | Dichlorofluoromethane | Dichlorofluoromethane | ND | ND | No | Yes | 0.00021407
120-83-2 | 2,4-Dichlorophenol | 2,4-dichlorophenol | ND | ND | No | Yes | 0.016267
78-87-5 | 1,2-Dichloropropane | 1,2-Dichloropropane | ND | ND | Yes | Yes | 0.00032906
10061-02-6 | trans-1,3-Dichloropropene | trans-1,3-Dichloropropene | ND | ND | No | Yes | 0.00023082
76-14-2 | Dichlorotetrafluoroethane | Freon 114 | ND | ND | No | Yes | 0.00049775
105-67-9 | 2,4-Dimethylphenol | 2,4-dimethylphenol | ND | ND | No | Yes | 0.016267
131-11-3 | Dimethyl phthalate | Dimethylphthalate | ND | ND | Yes | Yes | 0.016267
99-65-0 | 1,3-Dinitrobenzene | 1,3-Dinitrobenzene | ND | ND | No | Yes | 0.00067909
534-52-1 | 4,6-Dinitro-2-methylphenol | 4,6-dinitro-2-methylphenol | ND | ND | Yes | Yes | 0.016267
51-28-5 | 2,4-Dinitrophenol | 2,4-dinitrophenol | ND | ND | Yes | Yes | 0.016267
121-14-2 | 2,4-Dinitrotoluene | 2,4-dinitrotoluene | ND | ND | Yes | Yes | 0.016267
121-14-2 | 2,4-Dinitrotoluene | 2,4-Dinitrotoluene | ND | ND | Yes | Yes | 0.00067909
606-20-2 | 2,6-Dinitrotoluene | 2,6-dinitrotoluene | ND | ND | No | Yes | 0.016267
606-20-2 | 2,6-Dinitrotoluene | 2,6-Dinitrotoluene | ND | ND | No | Yes | 0.00067909
123-91-1 | 1,4-Dioxane | 1,4-Dioxane | ND | ND | Yes | Yes | 0.00032988
-- | Total dioxin/furan compounds | Total Dixoin | 1.6691e-13 | 4.3466e-11 | Yes | Yes | --
140-88-5 | Ethyl acrylate | Ethyl Acrylate | ND | ND | Yes | Yes | 0.0001666
100-41-4 | Ethylbenzene | Ethylbenzene | 1.6118e-09 | 4.1975e-07 | Yes | Yes | --
74-85-1 | Ethylene | Ethylene | 3.7792e-07 | 9.8418e-05 | No | Yes | --
117-81-7 | bis(2-Ethylhexyl)phthalate | bis(2-Ethylhexyl)phthalate | 2.9849e-08 | 7.7732e-06 | Yes | Yes | --
206-44-0 | Fluoranthene | Fluoranthene | ND | ND | Yes | Yes | 0.016267
206-44-0 | Fluoranthene | Fluoranthene | 1.8523e-10 | 4.8237e-08 | Yes | Yes | --
86-73-7 | Fluorene | Fluorene | ND | ND | Yes | No | 0.016267
86-73-7 | Fluorene | Fluorene | 4.4095e-10 | 1.1483e-07 | Yes | No | --
50-00-0 | Formaldehyde | Formaldehyde | 1.5505e-07 | 4.0376e-05 | Yes | Yes | --
76-13-1 | Freon 113 | Freon 113 | ND | ND | No | Yes | 0.00054696
35822-46-9 | 1,2,3,4,6,7,8-Heptachlorodibenzo-p-dioxin | 1,2,3,4,6,7,8-HpCDD | 4.0353e-14 | 1.0508e-11 | Yes | Yes | --
67562-39-4 | 1,2,3,4,6,7,8-Heptachlorodibenzofuran | 1,2,3,4,6,7,8-HpCDF | 2.9702e-14 | 7.7348e-12 | Yes | Yes | --
55673-89-7 | 1,2,3,4,7,8,9-Heptachlorodibenzofuran | 1234789-HPCDF | 0 | 0 | Yes | Yes | --
118-74-1 | Hexachlorobenzene | Hexachlorobenzene | ND | ND | Yes | Yes | 0.016267
87-68-3 | Hexachlorobutadiene | Hexachlorobutadiene | ND | ND | Yes | Yes | 0.00076003
87-68-3 | Hexachlorobutadiene | Hexachlorobutadiene | ND | ND | Yes | Yes | 0.016267
77-47-4 | Hexacholorocyclopentadiene | Hexachlorocyclopentadiene | ND | ND | Yes | Yes | 0.016267
39227-28-6 | 1,2,3,4,7,8-Hexachlorodibenzo-p-dioxin | 123478-HXCDD | 0 | 0 | Yes | Yes | --
57653-85-7 | 1,2,3,6,7,8-Hexachlorodibenzo-p-dioxin | 123678-HXCDD | 1.0152e-14 | 2.6436e-12 | Yes | Yes | --
19408-74-3 | 1,2,3,7,8,9-Hexachlorodibenzo-p-dioxin | 123789-HXCDD | 6.2003e-17 | 1.6147e-14 | Yes | Yes | --
70648-26-9 | 1,2,3,4,7,8-Hexachlorodibenzofuran | 1,2,3,4,7,8-HxCDF | 1.0018e-14 | 2.6089e-12 | Yes | Yes | --
57117-44-9 | 1,2,3,6,7,8-Hexachlorodibenzofuran | 1,2,3,6,7,8-HxCDF | 2.4305e-15 | 6.3294e-13 | Yes | Yes | --
72918-21-9 | 1,2,3,7,8,9-Hexachlorodibenzofuran | 123789-HXCDF | ND | ND | Yes | Yes | 3.6854e-09
60851-34-5 | 2,3,4,6,7,8-Hexachlorodibenzofuran | 2,3,4,6,7,8-HxCDF | 0 | 0 | Yes | Yes | --
67-72-1 | Hexachloroethane | Hexachloroethane | ND | ND | Yes | Yes | 0.016267
67-72-1 | Hexachloroethane | Hexachlorethane | ND | ND | Yes | Yes | 0.0005909
110-54-3 | Hexane | Hexane | 0 | 0 | Yes | Yes | --
110-54-3 | Hexane | n-Hexane | ND | ND | Yes | Yes | 0.071702
7647-01-0 | Hydrochloric acid | Hydrogen Chloride | ND | ND | Yes | Yes | 0.17452
74-90-8 | Hydrogen cyanide | Hydrogen Cyanide | 2.9718e-06 | 0.00077389 | Yes | Yes | --
7664-39-3 | Hydrogen fluoride | Hydrogen Fluoride | ND | ND | Yes | Yes | 0.18478
193-39-5 | Indeno[1,2,3-cd]pyrene | Indeno(1,2,3-cd)pyrene | ND | ND | Yes | Yes | 0.016267
193-39-5 | Indeno[1,2,3-cd]pyrene | Indeno(1,2,3-cd)pyrene | 8.1734e-11 | 2.1285e-08 | Yes | Yes | --
78-59-1 | Isophorone | Isophorone | ND | ND | Yes | No | 0.016267
7439-96-5 | Manganese | Manganese | ND | ND | Yes | Yes | 0.0045224
96-33-3 | Methyl acrylate | Methyl Acrylate | ND | ND | No | Yes | 0.00025072
74-95-3 | Methylene bromide | Dibromomethane | ND | ND | No | Yes | 0.00050619
75-09-2 | Methylene chloride | Methylene chloride | 0 | 0 | Yes | Yes | --
107-06-2 | 1,2-Dichloroethane | 1,2-Dichloroethane | 3.7574e-09 | 9.785e-07 | Yes | Yes | --
78-93-3 | Methyl ethyl ketone | 2-Butanone | 3.8245e-09 | 9.9597e-07 | Yes | Yes | --
74-88-4 | Methyl iodide | Methyl Iodide | ND | ND | Yes | Yes | 0.00035428
108-10-1 | Methyl isobutyl ketone | 4-Methyl-2-pentanone | ND | ND | Yes | Yes | 0.0003328
80-62-6 | Methyl methacrylate | Methyl Methacrylate | ND | ND | Yes | Yes | 0.00045815
91-57-6 | 2-Methylnaphthalene | 2-Methylnaphthalene | ND | ND | Yes | No | 0.016267
95-48-7 | 2-Methylphenol | 2-methylphenol | ND | ND | Yes | Yes | 0.016267
1634-04-4 | Methyl tert-butyl ether | Methyl t-Butyl Ether | 6.6837e-10 | 1.7406e-07 | Yes | Yes | --
91-20-3 | Naphthalene | Naphthalene | ND | ND | Yes | Yes | 0.016267
91-20-3 | Naphthalene | Naphthalene | 1.2712e-08 | 3.3104e-06 | Yes | Yes | --
7440-02-0 | Nickel | Nickel | ND | ND | Yes | Yes | 0.0090448
7697-37-2 | Nitric acid | Nitric Acid | 2.4809e-07 | 6.4606e-05 | No | Yes | --
100-01-6 | 4-Nitroaniline | 4-nitroaniline | ND | ND | No | Yes | 0.016267
98-95-3 | Nitrobenzene | Nitrobenzene | ND | ND | Yes | Yes | 0.016267
98-95-3 | Nitrobenzene | Nitrobenzene | ND | ND | Yes | Yes | 0.00067909
55-63-0 | Nitroglycerin | Nitroglycerine | ND | ND | No | Yes | 0.00069722
88-75-5 | 2-Nitrophenol | 2-nitrophenol | ND | ND | No | Yes | 0.016267
100-02-7 | 4-Nitrophenol | 4-nitrophenol | ND | ND | Yes | Yes | 0.016267
62-75-9 | n-Nitrosodimethylamine | N-nitrosodimethylamine | ND | ND | Yes | Yes | 0.016267
86-30-6 | n-Nitrosodiphenylamine | N-nitrosodiphenylamine(1) | ND | ND | Yes | Yes | 0.016267
621-64-7 | n-Nitroso-di-n-propylamine | N-nitroso-di-n-propylamine | ND | ND | No | Yes | 0.016267
3268-87-9 | 1,2,3,4,6,7,8,9-Octachlorodibenzo-p-dioxin | OCDD | 0 | 0 | Yes | Yes | --
39001-02-0 | 1,2,3,4,6,7,8,9-Octachlorodibenzofuran | OCDF | 0 | 0 | Yes | Yes | --
40321-76-4 | 1,2,3,7,8-Pentachlorodibenzo-p-dioxin | 12378-PECDD | 4.0302e-15 | 1.0495e-12 | Yes | Yes | --
57117-41-6 | 1,2,3,7,8-Pentachlorodibenzofuran | 12378-PECDF | 9.2835e-15 | 2.4176e-12 | Yes | Yes | --
57117-31-4 | 2,3,4,7,8-Pentachlorodibenzofuran | 23478-PECDF | 1.0192e-14 | 2.6541e-12 | Yes | Yes | --
87-86-5 | Pentachlorophenol | Pentachlorophenol | ND | ND | Yes | Yes | 0.016267
85-01-8 | Phenanthrene | Phenanthrene | ND | ND | Yes | Yes | 0.016267
85-01-8 | Phenanthrene | Phenanthrene | 4.6801e-10 | 1.2188e-07 | Yes | Yes | --
108-95-2 | Phenol | Phenol | ND | ND | Yes | Yes | 0.016267
123-38-6 | Propionaldehyde | Proprionaldehyde | ND | ND | Yes | Yes | 0.020524
115-07-1 | Propylene | Propene | 5.7379e-08 | 1.4942e-05 | No | Yes | --
115-07-1 | Propylene | Propylene | 5.2681e-08 | 1.3719e-05 | No | Yes | --
129-00-0 | Pyrene | Pyrene | ND | ND | Yes | No | 0.016267
129-00-0 | Pyrene | Pyrene | 4.0333e-10 | 1.0504e-07 | Yes | No | --
7782-49-2 | Selenium | Selenium | ND | ND | Yes | Yes | 0.0090448
7440-22-4 | Silver | Silver | ND | ND | No | Yes | 0.0022612
100-42-5 | Styrene | Styrene | 7.8944e-09 | 2.0558e-06 | Yes | Yes | --
7664-93-9 | Sulfuric acid | Sulfuric Acid | 3.5128e-07 | 9.148e-05 | No | Yes | --
1746-01-6 | 2,3,7,8-Tetrachlorodibenzo-p-dioxin | 2378-TCDD | 2.9159e-14 | 7.5935e-12 | Yes | Yes | --
51207-31-9 | 2,3,7,8-Tetrachlorodibenzofuran | 2,3,7,8-TCDF | 2.153e-14 | 5.6068e-12 | Yes | Yes | --
630-20-6 | 1,1,1,2-Tetrachloroethane | 1,1,1,2-Tetrachloroethane | ND | ND | No | Yes | 0.00034913
79-34-5 | 1,1,2,2-Tetrachloroethane | 1,1,2,2-Tetrachloroethane | ND | ND | Yes | Yes | 0.00034944
127-18-4 | Tetrachloroethylene | Tertrachloroethene | ND | ND | Yes | Yes | 0.00041391
7440-28-0 | Thallium | Thallium | 1.3748e-08 | 3.5803e-06 | No | Yes | --
108-88-3 | Toluene | Toluene | 1.536e-08 | 3.9999e-06 | Yes | Yes | --
120-82-1 | 1,2,4-Trichlorobenzene | 1,2,4-Trichlorobenzene | ND | ND | Yes | Yes | 0.00128
120-82-1 | 1,2,4-Trichlorobenzene | 1,2,4-trichlorobenzene | ND | ND | Yes | Yes | 0.016267
71-55-6 | 1,1,1-Trichloroethane | 1,1,1-Trichloroethane | ND | ND | Yes | Yes | 0.00044399
79-00-5 | 1,1,2-Trichloroethane | 1,1,2-Trichloroethane | ND | ND | Yes | Yes | 0.00027664
75-69-4 | Trichlorofluoromethane | Trichloroflouromethane | 0 | 0 | No | Yes | --
95-95-4 | 2,4,5-Trichlorophenol | 2,4,5-trichlorophenol | ND | ND | Yes | Yes | 0.016267
88-06-2 | 2,4,6-Trichlorophenol | 2,4,6-trichlorophenol | ND | ND | Yes | Yes | 0.016267
96-18-4 | 1,2,3-Trichloropropane | 1,2,3-Trichloropropane | ND | ND | No | Yes | 0.00049065
95-63-6 | 1,2,4-Trimethylbenzene | 1,2,4-Trimethylbenzene | 0 | 0 | No | Yes | --
540-84-1 | 2,2,4-Trimethylpentane | Isooctane | ND | ND | Yes | No | 0.00019008
7440-62-2 | Vanadium | Vanadium | ND | ND | No | Yes | 0.0045224
108-05-4 | Vinyl acetate | Vinyl Acetate | ND | ND | Yes | Yes | 0.00039395
75-01-4 | Vinyl chloride | Vinyl Chloride | ND | ND | Yes | Yes | 0.00026
75-35-4 | Vinylidene chloride | 1,1-Dichlorethene | ND | ND | Yes | Yes | 0.00020584
106-42-3, 108-38-3 | m-Xylene, p-Xylene | m&p-Xylene | 0 | 0 | Yes | Yes | --
95-47-6 | o-Xylene | o-Xylene | 0 | 0 | Yes | Yes | --
7440-66-6 | Zinc | Zinc | 2.3808e-06 | 0.00062001 | No | Yes | --
Other Pollutants
67-64-1 | Acetone | Acetone | 7.0884e-08 | 1.8459e-05 | No | No | --
67-64-1 | Acetone | Acetone | 0 | 0 | No | No | --
74-86-2 | Acetylene | Acetylene | 5.4359e-08 | 1.4156e-05 | No | No | --
35572-78-2 | 2-Amino-4,6-dinitrotoluene | 2-Amino-4,6-Dinitrotoluene | ND | ND | No | No | 0.00067909
19406-51-0 | 4-Amino-2,6-dinitrotoluene | 4-Amino-2,6-Dinitrotoluene | ND | ND | No | No | 0.00067909
100-52-7 | Benzaldehyde | Benzaldehyde | ND | ND | No | No | 0.020524
4748-78-1 | Benzaldehyde, 4-ethyl- | Benzaldehyde, 4-ethyl- | 2.2533e-08 | 5.8679e-06 | No | No | --
100-51-6 | Benzyl alcohol | Benzyl alcohol | ND | ND | No | No | 0.016267
108-86-1 | Bromobenzene | Bromobenzene | ND | ND | No | No | 0.0003919
106-97-8 | Butane | Butane | ND | ND | No | No | 0.053191
106-98-9 | 1-Butene | 1-Butene | ND | ND | No | No | 0.046684
590-18-1 | cis-Butene | cis-butene | ND | ND | No | No | 0.046692
624-64-6 | trans-Butene | trans-Butene | ND | ND | No | No | 0.046692
107-00-6 | 1-Butyne | 1-Butyne | ND | ND | No | No | 0.045003
503-17-3 | 2-Butyne | 2-Butyne | ND | ND | No | No | 0.045003
7440-70-2 | Calcium | Calcium | 2.3022e-07 | 5.9954e-05 | No | No | --
7440-44-0 | Elemental carbon | Elemental Carbon | 9.2963e-07 | 0.00024209 | No | No | --
7440-44-0 | Organic carbon | Organic Carbon | 3.9123e-06 | 0.0010188 | No | No | --
7440-44-0 | Total carbon | Total Carbon | 4.842e-06 | 0.0012609 | No | No | --
39638-32-9 | bis(2-Chloroisopropyl)ether | Bis(2-chloroisopropyl)ether | ND | ND | No | No | 0.016267
59-50-7 | 4-Chloro-3-methylphenol | 4-chloro-3-methylphenol | ND | ND | No | No | 0.016267
124-48-1 | Dibromochloromethane | Dibromochloromethane | ND | ND | No | No | 0.00051969
156-59-2 | cis-1,2-Dichloroethene | cis-1,2-Dichloroethene | ND | ND | No | No | 0.00020164
156-60-5 | trans-1,2-Dichloroethene | trans-1,2,-Dichloroethene | ND | ND | No | No | 0.00048422
10061-01-5 | cis-1,3-Dichloropropylene | cis-1,3-Dichloropropene | ND | ND | No | No | 0.00027698
84-66-2 | Diethylphthalate | Diethylphthalate | ND | ND | No | No | 0.016267
5779-94-2 | 2,5-Dimethylbenzaldehyde | 2,5-Dimethylbenzaldehyde | ND | ND | No | No | 0.020524
117-84-0 | Di-n-octylphthalate | Di-n-octylphthalate | ND | ND | No | No | 0.016267
112-95-8 | Eicosane | Eicosane | 1.3638e-08 | 3.5516e-06 | No | No | --
74-84-0 | Ethane | Ethane | 1.9515e-07 | 5.0821e-05 | No | No | --
64-17-5 | Ethanol | Ethanol | 0 | 0 | No | No | --
141-78-6 | Ethyl acetate | Ethyl Acetate | ND | ND | No | No | 0.0013562
97-63-2 | Ethyl methacrylate | Ethyl Methacrylate | ND | ND | No | No | 0.00033238
622-96-8 | 4-Ethyltoluene | 4-Ethyltoluene | 0 | 0 | No | No | --
142-82-5 | n-Heptane | Heptane | ND | ND | No | No | 0.00020842
66-25-1 | Hexaldehyde | Hexaldehyde | ND | ND | No | No | 0.020524
591-78-6 | 2-Hexanone | 2-Hexanone | ND | ND | No | No | 0.0002912
2691-41-0 | HMX | HMX | ND | ND | No | No | 0.0067909
10035-10-6 | Hydrogen bromide | Hydrogen Bromide | ND | ND | No | No | 0.17452
123-92-2 | Isoamyl acetate | 1-Butanol, 3-methyl-, acetate | 1.6327e-08 | 4.2519e-06 | No | No | --
75-28-5 | Isobutane | Isobutane | ND | ND | No | No | 0.048356
626-19-7 | Isophthalaldehyde | Isophthalaldehyde | 3.5578e-08 | 9.2651e-06 | No | No | --
590-86-3 | Isovaleraldehyde | Isovaleraldehyde | ND | ND | No | No | 0.020524
7439-95-4 | Magnesium | Magnesium | 4.195e-07 | 0.00010925 | No | No | --
98-83-9 | alpha-Methylstyrene | Alpha Methyl Styrene | ND | ND | No | No | 0.00029498
479-45-8 | Methyl-2,4,6-trinitrophenylnitramine | Tetryl | ND | ND | No | No | 0.00067909
88-74-7 | 2-Nitroaniline | 2-nitroaniline | ND | ND | No | No | 0.016267
99-09-2 | 3-Nitroaniline | 3-nitroaniline | ND | ND | No | No | 0.016267
88-72-2 | 2-Nitrotoluene | 2-Nitrotoluene | ND | ND | No | No | 0.00067909
99-08-1 | 3-Nitrotoluene | 3-Nitrotoluene | ND | ND | No | No | 0.00067909
99-99-0 | 4-Nitrotoluene | 4-Nitrotoluene | ND | ND | No | No | 0.00067909
111-65-9 | Octane | Octane | ND | ND | No | No | 0.00019008
109-66-0 | Pentane | Pentane | 8.2059e-10 | 2.1369e-07 | No | No | --
109-66-0 | Pentane | n-Pentane | ND | ND | No | No | 0.060029
7664-38-2 | Phosphoric acid | Phosphoric Acid | ND | ND | No | No | 0.17452
74-98-6 | Propane | Propane | ND | ND | No | No | 0.036691
115-11-7 | 1-Propene, 2-methyl | Isobutylene | ND | ND | No | No | 0.046675
74-99-7 | Propyne | Propyne | ND | ND | No | No | 0.032531
121-82-4 | RDX | RDX | ND | ND | No | No | 0.00067909
1334-78-7 | o,m,p-Tolualdehyde | o,m,p-Tolualdehyde | ND | ND | No | No | 0.061571
25323-89-1 | Trichloroethane | Trichloroethane | ND | ND | No | No | 0.00029797
108-67-8 | 1,3,5-Trimethylbenzene | 1,3,5-Trimethylbenzene | 0 | 0 | No | No | --
99-35-4 | 1,3,5-Trinitrobenzene | 1,3,5-Trinitrobenzene | ND | ND | No | No | 0.00067909
118-96-7 | 2,4,6-Trinitrotoluene | 2,4,6-Trinitrotoluene | ND | ND | No | No | 0.00067909
110-62-3 | Valeraldehyde | Valeraldehyde | ND | ND | No | No | 0.020524
Note 1: Green shading indicates a compound that was sampled using two methods. In such instances, the emission factor reported in AP-42 was calculated using the more accurate test method, as indicated by the shading.
Note 2: Orange shading indicates tentatively identified compounds (TICs).
Note 3: Pink shading indicates a compound that was excluded from AP-42 due to its presence in the lab blank at levels greater than 50% of test levels.
Note 4: lb/lb NEW factors may be slightly different for the same lb/item value due to rounding.
Note 5: Red shading indicates test results that did not uniquely identify a compound. These compounds were not included in AP-42.
Note 6: Yellow shading indicates that a test or analytical method was selected that would not typically be chosen for the compound.
Note 7: Gray shading indicates that the test notation for this compound differs from that used in other tests. Name from the "renamed compounds" list was adopted.
Note 8: Blue shading indicates poor precision. These compounds had a relative percent difference greater than 100% within a test.
Note 9: Purple highlights indicate compounds present in the lab blank at levels between 20 and 50 % of test values. These compounds need to be annotated in AP-42 in the following way. These emission factors are assumed to be biased high.
Note 10: Teal highlights indicate the holding time allowance was significantly exceeded prior to analysis. Compounds will be downgraded in AP-42.
Note on F10: Kaobrian: lead compounds
Note on F18: Kaobrian: POM
Note on F19: Kaobrian: POM
Note on F20: Kaobrian: POM
Note on F21: Kaobrian: POM
Note on F31: Kaobrian: POM
Note on F32: Kaobrian: POM
Note on F33: Kaobrian: "antimony compounds"
Note on F34: Kaobrian: Arsenic Compounds
Note on F37: Kaobrian: POM
Note on F38: Kaobrian: POM
Note on F39: Kaobrian: POM
Note on F40: Kaobrian: POM
Note on F41: Kaobrian: POM
Note on F42: Kaobrian: POM
Note on F43: Kaobrian: POM
Note on F44: Kaobrian: POM
Note on F45: Kaobrian: POM
Note on F46: Kaobrian: POM
Note on F47: Kaobrian: POM
Note on F49: Kaobrian: Beryllium Compounds
Note on F52: Kaobrian: POM
Note on F56: Kaobrian: POM
Note on F58: Kaobrian: cadmium compounds
Note on F69: Kaobrian: POM
Note on F71: Kaobrian: POM
Note on F72: Kaobrian: Chromium compounds
Note on F73: Kaobrian: Chromium Compounds
Note on F74: Kaobrian: POM
Note on F75: Kaobrian: POM
Note on F76: Kaobrian: cobalt compounds
Note on F81: Kaobrian: cyanide compounds
Note on F82: Kaobrian: POM
Note on F83: Kaobrian: POM
Note on F116: Kaobrian: POM
Note on F117: Kaobrian: POM
Note on F118: Kaobrian: POM
Note on F119: Kaobrian: POM
Note on F141: Kaobrian: cyanide compounds
Note on F143: Kaobrian: POM
Note on F144: Kaobrian: POM
Note on F146: Kaobrian: manganese compounds
Note on F155: Kaobrian: POM
Note on F160: Kaobrian: Nickel compounds
Note on F169: Kaobrian: POM
Note on F177: Kaobrian: POM
Note on F178: Kaobrian: POM
Note on F183: Kaobrian: POM
Note on F184: Kaobrian: POM
Note on F185: Kaobrian: Selenium Compounds
Note on F189: Kaobrian: POM
Note on G52: Kaobrian: polybrominated biphenyl
Note on G70: Kaobrian: chlorophenol
Note on G73: Kaobrian: Chromium Compounds

## Sum All EFs
Summary of CEM, FTIR, Acid Gases, Cyanide, Particulates, Metals, Aldehyde, VOC, Hydrocarbon, SVOC, TO-13-PAH, Dioxin-Furan, Energetic, and Emission Factors
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
Pollutant | Emission Factor |  | Detection Limit Based Concentration (mg/m3) | Relative Percent Difference
| (lb/item) | (lb/lb NEW)
CEM
NOx | 1.7461e-05 | 0.0045454 | -- | 0%
SO2 | ND | ND | 0.53248 | 0%
CO | 0.0016183 | 0.42131 | -- | 0%
CO2 | 0.00097761 | 0.2545 | -- | 0%
FTIR
NH3 | 3.7829e-05 | 0.0098479 | -- | 0%
Acid Gases
Hydrogen Fluoride | ND | ND | 0.18478 | ND
Hydrogen Chloride | ND | ND | 0.17452 | ND
Hydrogen Bromide | ND | ND | 0.17452 | ND
Nitric Acid | 2.4809e-07 | 6.4606e-05 | -- | 56.64%
Phosphoric Acid | ND | ND | 0.17452 | ND
Sulfuric Acid | 3.5128e-07 | 9.148e-05 | -- | -77.3%
Cyanide
Particulate Cyanide | 1.2383e-06 | 0.00032247 | -- | 64.86%
Hydrogen Cyanide | 2.9718e-06 | 0.00077389 | -- | -162.3%
Particulates
TSP | 7.0157e-05 | 0.01827 | -- | 1.293%
PM10 | 6.7046e-05 | 0.01746 | -- | 1.509%
PM2.5 | 5.1225e-05 | 0.01334 | -- | 0.2408%
Metals
Aluminum | 1.1024e-07 | 2.8708e-05 | -- | -2.377%
Antimony | 1.325e-06 | 0.00034504 | -- | 1.043%
Arsenic | ND | ND | 0.0090448 | ND
Barium | 4.6637e-07 | 0.00012145 | -- | -2.53%
Beryllium | ND | ND | 0.0045224 | ND
Cadmium | ND | ND | 0.0045224 | ND
Calcium | 2.3022e-07 | 5.9954e-05 | -- | 8.39%
Chromium | ND | ND | 0.0022612 | ND
Cobalt | ND | ND | 0.0045224 | ND
Copper | 2.0743e-05 | 0.0054017 | -- | -2.71%
Lead | 2.7632e-06 | 0.00071959 | -- | 2.831%
Magnesium | 4.195e-07 | 0.00010925 | -- | 9.09%
Manganese | ND | ND | 0.0045224 | ND
Nickel | ND | ND | 0.0090448 | ND
Selenium | ND | ND | 0.0090448 | ND
Silver | ND | ND | 0.0022612 | ND
Thallium | 1.3748e-08 | 3.5803e-06 | -- | -79.9%
Vanadium | ND | ND | 0.0045224 | ND
Zinc | 2.3808e-06 | 0.00062001 | -- | -4.79%
Aldehyde
Formaldehyde | 1.5505e-07 | 4.0376e-05 | -- | 1.044%
Acetaldehyde | 1.231e-07 | 3.2058e-05 | -- | -31.81%
Acetone | 7.0884e-08 | 1.8459e-05 | -- | 103.8%
Acrolein | ND | ND | 0.01334 | ND
Proprionaldehyde | ND | ND | 0.020524 | ND
Crotonaldehyde | ND | ND | 0.020524 | ND
Butyraldehyde | ND | ND | 0.020524 | ND
Benzaldehyde | ND | ND | 0.020524 | ND
Isovaleraldehyde | ND | ND | 0.020524 | ND
Valeraldehyde | ND | ND | 0.020524 | ND
o,m,p-Tolualdehyde | ND | ND | 0.061571 | ND
Hexaldehyde | ND | ND | 0.020524 | ND
2,5-Dimethylbenzaldehyde | ND | ND | 0.020524 | ND
VOC
Propene | 5.7379e-08 | 1.4942e-05 | -- | 22.22%
Dichlorodifluoromethane | 0 | 0 | -- | 0%
Chlorodifluoromethane | ND | ND | 0.0002518 | ND
Freon 114 | ND | ND | 0.00049775 | ND
Chloromethane | 2.2995e-09 | 5.9884e-07 | -- | 0%
Vinyl Chloride | ND | ND | 0.00026 | ND
1,3-Butadiene | 1.0245e-08 | 2.668e-06 | -- | 40%
Bromomethane | ND | ND | 0.00035546 | ND
Chloroethane | ND | ND | 0.00061733 | ND
Dichlorofluoromethane | ND | ND | 0.00021407 | ND
Trichloroflouromethane | 0 | 0 | -- | 0%
Pentane | 8.2059e-10 | 2.1369e-07 | -- | 173.8%
Acrolein | 2.6472e-09 | 6.8937e-07 | -- | 121.3%
1,1-Dichlorethene | ND | ND | 0.00020584 | ND
Freon 113 | ND | ND | 0.00054696 | ND
Acetone | 0 | 0 | -- | 0%
Methyl Iodide | ND | ND | 0.00035428 | ND
Carbon Disulfide | ND | ND | 0.00031616 | ND
Acetonitrile | 2.3496e-07 | 6.1187e-05 | -- | 9.272%
3-Chloropropene | ND | ND | 0.00089142 | ND
Methylene chloride | 0 | 0 | -- | -7.407%
tert-Butyl Alcohol | ND | ND | 0.00040084 | ND
Acrylonitrile | 2.0116e-08 | 5.2384e-06 | -- | 0%
trans-1,2,-Dichloroethene | ND | ND | 0.00048422 | ND
Methyl t-Butyl Ether | 6.6837e-10 | 1.7406e-07 | -- | 163.6%
Hexane | 0 | 0 | -- | 40%
1,1-Dichloroethane | ND | ND | 0.00037066 | ND
Vinyl Acetate | ND | ND | 0.00039395 | ND
cis-1,2-Dichloroethene | ND | ND | 0.00020164 | ND
2-Butanone | 3.8245e-09 | 9.9597e-07 | -- | 57.14%
Ethyl Acetate | ND | ND | 0.0013562 | ND
Methyl Acrylate | ND | ND | 0.00025072 | ND
Chloroform | ND | ND | 0.00039603 | ND
1,1,1-Trichloroethane | ND | ND | 0.00044399 | ND
Carbontetrachloride | ND | ND | 0.00032032 | ND
1,2-Dichloroethane | 3.7574e-09 | 9.785e-07 | -- | 0%
Benzene | 1.9242e-07 | 5.011e-05 | -- | 3.077%
Isooctane | ND | ND | 0.00019008 | ND
Heptane | ND | ND | 0.00020842 | ND
Trichloroethane | ND | ND | 0.00029797 | ND
Ethyl Acrylate | ND | ND | 0.0001666 | ND
1,2-Dichloropropane | ND | ND | 0.00032906 | ND
Methyl Methacrylate | ND | ND | 0.00045815 | ND
Dibromomethane | ND | ND | 0.00050619 | ND
1,4-Dioxane | ND | ND | 0.00032988 | ND
Bromodichloromethane | ND | ND | 0.00040892 | ND
cis-1,3-Dichloropropene | ND | ND | 0.00027698 | ND
4-Methyl-2-pentanone | ND | ND | 0.0003328 | ND
Toluene | 1.536e-08 | 3.9999e-06 | -- | 40%
Octane | ND | ND | 0.00019008 | ND
trans-1,3-Dichloropropene | ND | ND | 0.00023082 | ND
Ethyl Methacrylate | ND | ND | 0.00033238 | ND
1,1,2-Trichloroethane | ND | ND | 0.00027664 | ND
Tertrachloroethene | ND | ND | 0.00041391 | ND
2-Hexanone | ND | ND | 0.0002912 | ND
Dibromochloromethane | ND | ND | 0.00051969 | ND
1,2-Dibromoethane | ND | ND | 0.00039104 | ND
Chlorobenzene | 9.6582e-10 | 2.5152e-07 | -- | -155.6%
1,1,1,2-Tetrachloroethane | ND | ND | 0.00034913 | ND
Ethylbenzene | 1.6118e-09 | 4.1975e-07 | -- | 0%
m&p-Xylene | 0 | 0 | -- | -66.67%
o-Xylene | 0 | 0 | -- | 0%
Styrene | 7.8944e-09 | 2.0558e-06 | -- | 0%
Bromoform | ND | ND | 0.00052568 | ND
Cumene | ND | ND | 0.0002 | ND
1,1,2,2-Tetrachloroethane | ND | ND | 0.00034944 | ND
1,2,3-Trichloropropane | ND | ND | 0.00049065 | ND
Bromobenzene | ND | ND | 0.0003919 | ND
4-Ethyltoluene | 0 | 0 | -- | 0%
1,3,5-Trimethylbenzene | 0 | 0 | -- | -169.2%
Alpha Methyl Styrene | ND | ND | 0.00029498 | ND
1,2,4-Trimethylbenzene | 0 | 0 | -- | -46.15%
1,3-Dichlorobenzene | ND | ND | 0.00042806 | ND
1,4-Dichlorobenzene | ND | ND | 0.00042806 | ND
Benzylchloride | ND | ND | 0.00026416 | ND
1,2-Dichlorobenzene | ND | ND | 0.00048922 | ND
Hexachlorethane | ND | ND | 0.0005909 | ND
1,2,4-Trichlorobenzene | ND | ND | 0.00128 | ND
Hexachlorobutadiene | ND | ND | 0.00076003 | ND
VOC TIC
Ethanol | 0 | 0 | -- | 0%
Hydrocarbons
Methane | 6.2094e-06 | 0.001617 | -- | 2.857%
1-Butene | ND | ND | 0.046684 | ND
1-Butyne | ND | ND | 0.045003 | ND
2-Butyne | ND | ND | 0.045003 | ND
Acetylene | 5.4359e-08 | 1.4156e-05 | -- | -3.636%
1,3-Butadiene | ND | ND | 0.045003 | ND
Butane | ND | ND | 0.053191 | ND
cis-butene | ND | ND | 0.046692 | ND
Ethane | 1.9515e-07 | 5.0821e-05 | -- | 3.509%
Ethylene | 3.7792e-07 | 9.8418e-05 | -- | 5.07%
n-Hexane | ND | ND | 0.071702 | ND
Isobutane | ND | ND | 0.048356 | ND
Isobutylene | ND | ND | 0.046675 | ND
n-Pentane | ND | ND | 0.060029 | ND
Propane | ND | ND | 0.036691 | ND
Propylene | 5.2681e-08 | 1.3719e-05 | -- | 78.79%
Propyne | ND | ND | 0.032531 | ND
trans-Butene | ND | ND | 0.046692 | ND
SVOC
N-nitrosodimethylamine | ND | ND | 0.016267 | ND
Bis(2-chloroethyl)ether | ND | ND | 0.016267 | ND
Phenol | ND | ND | 0.016267 | ND
2-chlorophenol | ND | ND | 0.016267 | ND
1,3-dichlorobenzene | ND | ND | 0.016267 | ND
1,4-dichlorobenzene | ND | ND | 0.016267 | ND
1,2-dichlorobenzene | ND | ND | 0.016267 | ND
Benzyl alcohol | ND | ND | 0.016267 | ND
Bis(2-chloroisopropyl)ether | ND | ND | 0.016267 | ND
2-methylphenol | ND | ND | 0.016267 | ND
Hexachloroethane | ND | ND | 0.016267 | ND
N-nitroso-di-n-propylamine | ND | ND | 0.016267 | ND
4-methylphenol | ND | ND | 0.016267 | ND
Nitrobenzene | ND | ND | 0.016267 | ND
Isophorone | ND | ND | 0.016267 | ND
2-nitrophenol | ND | ND | 0.016267 | ND
2,4-dimethylphenol | ND | ND | 0.016267 | ND
Bis(2-chloroethoxy)methane | ND | ND | 0.016267 | ND
2,4-dichlorophenol | ND | ND | 0.016267 | ND
1,2,4-trichlorobenzene | ND | ND | 0.016267 | ND
Naphthalene | ND | ND | 0.016267 | ND
4-chloroaniline | ND | ND | 0.016267 | ND
Hexachlorobutadiene | ND | ND | 0.016267 | ND
4-chloro-3-methylphenol | ND | ND | 0.016267 | ND
2-Methylnaphthalene | ND | ND | 0.016267 | ND
Hexachlorocyclopentadiene | ND | ND | 0.016267 | ND
2,4,6-trichlorophenol | ND | ND | 0.016267 | ND
2,4,5-trichlorophenol | ND | ND | 0.016267 | ND
2-chloronaphthalene | ND | ND | 0.016267 | ND
2-nitroaniline | ND | ND | 0.016267 | ND
Acenaphthylene | ND | ND | 0.016267 | ND
Dimethylphthalate | ND | ND | 0.016267 | ND
2,6-dinitrotoluene | ND | ND | 0.016267 | ND
Acenaphthene | ND | ND | 0.016267 | ND
3-nitroaniline | ND | ND | 0.016267 | ND
2,4-dinitrophenol | ND | ND | 0.016267 | ND
Dibenzofuran | ND | ND | 0.016267 | ND
2,4-dinitrotoluene | ND | ND | 0.016267 | ND
4-nitrophenol | ND | ND | 0.016267 | ND
Fluorene | ND | ND | 0.016267 | ND
4-chlorophenyl-phenylether | ND | ND | 0.016267 | ND
Diethylphthalate | ND | ND | 0.016267 | ND
4-nitroaniline | ND | ND | 0.016267 | ND
4,6-dinitro-2-methylphenol | ND | ND | 0.016267 | ND
N-nitrosodiphenylamine(1) | ND | ND | 0.016267 | ND
4-bromophenyl-phenylether | ND | ND | 0.016267 | ND
Hexachlorobenzene | ND | ND | 0.016267 | ND
Pentachlorophenol | ND | ND | 0.016267 | ND
Phenanthrene | ND | ND | 0.016267 | ND
Anthracene | ND | ND | 0.016267 | ND
Di-n-butylphthalate | ND | ND | 0.016267 | ND
Fluoranthene | ND | ND | 0.016267 | ND
Pyrene | ND | ND | 0.016267 | ND
Butylbenzylphthalate | ND | ND | 0.016267 | ND
Benzo(a)anthracene | ND | ND | 0.016267 | ND
Chrysene | ND | ND | 0.016267 | ND
bis(2-Ethylhexyl)phthalate | 2.9849e-08 | 7.7732e-06 | -- | 98.15%
Di-n-octylphthalate | ND | ND | 0.016267 | ND
Benzo(b)fluoranthene | ND | ND | 0.016267 | ND
Benzo(k)fluoranthene | ND | ND | 0.016267 | ND
Benzo(a)pyrene | ND | ND | 0.016267 | ND
Indeno(1,2,3-cd)pyrene | ND | ND | 0.016267 | ND
Dibenz(a,h)anthracene | ND | ND | 0.016267 | ND
Benzo(g,h,i)perylene | ND | ND | 0.016267 | ND
SVOC TIC
1-Butanol, 3-methyl-, acetate | 1.6327e-08 | 4.2519e-06 | -- | 0%
Aectophenone | 3.2613e-08 | 8.493e-06 | -- | 0%
Isophthalaldehyde | 3.5578e-08 | 9.2651e-06 | -- | 0%
Benzaldehyde, 4-ethyl- | 2.2533e-08 | 5.8679e-06 | -- | 0%
Eicosane | 1.3638e-08 | 3.5516e-06 | -- | 0%
TO-13-PAH
Naphthalene | 1.2712e-08 | 3.3104e-06 | -- | 6.646%
Acenaphthylene | 1.2012e-09 | 3.1281e-07 | -- | 7.556%
Acenaphthene | 1.7354e-10 | 4.5193e-08 | -- | 6.858%
Fluorene | 4.4095e-10 | 1.1483e-07 | -- | 3.267%
Phenanthrene | 4.6801e-10 | 1.2188e-07 | -- | 19.8%
Anthracene | 9.1271e-11 | 2.3769e-08 | -- | 60.08%
Fluoranthene | 1.8523e-10 | 4.8237e-08 | -- | 132%
Pyrene | 4.0333e-10 | 1.0504e-07 | -- | 153.1%
Benzo(a)anthracene | ND | ND | 8.1333e-05 | ND
Chrysene | ND | ND | 8.1333e-05 | ND
Benzo(b)fluoranthene | 1.2108e-10 | 3.1531e-08 | -- | 74.44%
Benzo(k)fluoranthene | ND | ND | 8.1333e-05 | ND
Benzo(e)pyrene | 1.2256e-10 | 3.1917e-08 | -- | 75.96%
Benzo(a)pyrene | 1.196e-10 | 3.1145e-08 | -- | 72.89%
Indeno(1,2,3-cd)pyrene | 8.1734e-11 | 2.1285e-08 | -- | 162.8%
Dibenz(a,h)anthracene | ND | ND | 1.6267e-05 | ND
Benzo(g,h,i)perylene | 2.0033e-10 | 5.2169e-08 | -- | 184.8%
Dioxin-Furan
2378-TCDD | 2.9159e-14 | 7.5935e-12 | -- | 11.29%
12378-PECDD | 4.0302e-15 | 1.0495e-12 | -- | 105.2%
123478-HXCDD | 0 | 0 | -- | 113.1%
123678-HXCDD | 1.0152e-14 | 2.6436e-12 | -- | 105.8%
123789-HXCDD | 6.2003e-17 | 1.6147e-14 | -- | 99.66%
1,2,3,4,6,7,8-HpCDD | 4.0353e-14 | 1.0508e-11 | -- | 108.1%
OCDD | 0 | 0 | -- | 85.21%
2,3,7,8-TCDF | 2.153e-14 | 5.6068e-12 | -- | 103.2%
12378-PECDF | 9.2835e-15 | 2.4176e-12 | -- | 105.9%
23478-PECDF | 1.0192e-14 | 2.6541e-12 | -- | 117.4%
1,2,3,4,7,8-HxCDF | 1.0018e-14 | 2.6089e-12 | -- | 102.3%
1,2,3,6,7,8-HxCDF | 2.4305e-15 | 6.3294e-13 | -- | 94.24%
123789-HXCDF | ND | ND | 3.6854e-09 | ND
2,3,4,6,7,8-HxCDF | 0 | 0 | -- | 98.66%
1,2,3,4,6,7,8-HpCDF | 2.9702e-14 | 7.7348e-12 | -- | 106.9%
1234789-HPCDF | 0 | 0 | -- | -0.009566%
OCDF | 0 | 0 | -- | 25.69%
Total Dixoin | 1.6691e-13 | 4.3466e-11 | -- | --
Energetic
Nitrobenzene | ND | ND | 0.00067909 | 0%
2-Nitrotoluene | ND | ND | 0.00067909 | 0%
3-Nitrotoluene | ND | ND | 0.00067909 | 0%
4-Nitrotoluene | ND | ND | 0.00067909 | 0%
Nitroglycerine | ND | ND | 0.00069722 | 0%
1,3-Dinitrobenzene | ND | ND | 0.00067909 | 0%
2,6-Dinitrotoluene | ND | ND | 0.00067909 | 0%
2,4-Dinitrotoluene | ND | ND | 0.00067909 | 0%
1,3,5-Trinitrobenzene | ND | ND | 0.00067909 | 0%
2,4,6-Trinitrotoluene | ND | ND | 0.00067909 | 0%
RDX | ND | ND | 0.00067909 | 0%
4-Amino-2,6-Dinitrotoluene | ND | ND | 0.00067909 | 0%
2-Amino-4,6-Dinitrotoluene | ND | ND | 0.00067909 | 0%
Tetryl | ND | ND | 0.00067909 | 0%
HMX | ND | ND | 0.0067909 | 0%
Hexavalent Chromium
Hexavalent Chromium | ND | ND | 0.00060545 | ND
Carbon
Organic Carbon | 3.9123e-06 | 0.0010188 | -- | 13.83%
Elemental Carbon | 9.2963e-07 | 0.00024209 | -- | 15.06%
Total Carbon | 4.842e-06 | 0.0012609 | -- | --

## CEM & FTIR
CEM and FTIR Data and Emission Factor Calculations
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
Test Run Number | 1 | 2 | 3 | 4 | Background
Number of Items Tested | 20 | 20 | 20 | 20 | 0
Net Explosive Weight for All Items (lb) | 0.0768 | 0.0768 | 0.076852 | 0.076852 | NA
Initial Plume Volume (m3) | 14.716 | 14.716 | 14.716 | 14.716 | 14.716
Dilution Correction Factor | NA | NA | NA | NA | NA
Is Sample 1 Active? | Yes | Yes | Yes | Yes | Yes
Is Sample 2 Active? | No | No | No | No | No
CEM-Measured Compounds | [9 empty] | CEM-Measured Compounds
Pollutant | Sample ID | Sample Speciation (ppmv) |  |  |  | Detection Limits (ppmv) |  |  |  | Pollutant | Run Number | Molecular Weight (g/gmol) | Sample Speciation with Detection Limits (SSDL) (ppmv)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Quantity of Pollutant Emitted (QPE) (lb)c |  | Emission Factor (lb/item)d |  | Emission Factor (lb/lb NEW)e |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)f |  | Minimum Detection Limit Based Concentration (mg/m3)
|  | Run 1 | Run 2 | Run 3 | Run 4 | Run 1 | Run 2 | Run 3 | Run 4 |  |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
NOx | Sample 1 | 5.2 | 4.8 | 6.3 | 6.2 | -- | -- | -- | -- | NOx | 1 | 46 | 5.2 | ND | 9.9507 | ND | 0.00032283 | ND | 1.6142e-05 | ND | 0.0042036 | ND | 1.7461e-05 | 0.0045454 | -- | Not Applicable | --
| Sample 2 | [9 empty] | 2 |  | 4.8 | ND | 9.1853 | ND | 0.000298 | ND | 1.49e-05 | ND | 0.0038802 | ND |  |  | -- | Not Applicable | bar: Sample Speciation (ppmv) Run 1 to Detection Limits (ppmv) Run 4
| Background | [9 empty] | 3 |  | 6.3 | ND | 12.0557 | ND | 0.00039112 | ND | 1.9556e-05 | ND | 0.0050893 | ND |  |  | -- | Not Applicable | bar: Sample Speciation (ppmv) Run 1 to Detection Limits (ppmv) Run 4
[11 empty] | 4 |  | 6.2 | ND | 11.8643 | ND | 0.00038492 | ND | 1.9246e-05 | ND | 0.0050086 | ND |  |  | -- | Not Applicable
SO2 | Sample 1 | ND | ND | ND | ND | 0.2 | 0.2 | 0.2 | 0.2 | SO2 | 1 | 64 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.53248 | Not Applicable | 0.53248
| Sample 2 | [9 empty] | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.53248 | Not Applicable | bar: Sample Speciation (ppmv) Run 1 to Detection Limits (ppmv) Run 4
| Background | ND | ND | ND | ND | 0.2 | 0.2 | 0.2 | 0.2 |  | 3 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.53248 | Not Applicable
[11 empty] | 4 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.53248 | Not Applicable
CO | Sample 1 | 877 | 883 | 840 | 826 | -- | -- | -- | -- | CO | 1 | 28 | 877 | ND | 1021.5296 | ND | 0.033142 | ND | 0.0016571 | ND | 0.43153 | ND | 0.0016183 | 0.42131 | -- | Not Applicable | --
| Sample 2 | [9 empty] | 2 |  | 883 | ND | 1028.5184 | ND | 0.033368 | ND | 0.0016684 | ND | 0.43448 | ND |  |  | -- | Not Applicable | bar: Sample Speciation (ppmv) Run 1 to Detection Limits (ppmv) Run 4
| Background | ND | ND | ND | ND | 40 | 40 | 40 | 40 |  | 3 |  | 840 | ND | 978.432 | ND | 0.031743 | ND | 0.0015872 | ND | 0.41305 | ND |  |  | -- | Not Applicable
[11 empty] | 4 |  | 826 | ND | 962.1248 | ND | 0.031214 | ND | 0.0015607 | ND | 0.40616 | ND |  |  | -- | Not Applicable
CO2 | Sample 1 | 333 | 337 | 324 | 323 | -- | -- | -- | -- | CO2 | 1 | 44 | 333 | ND | 609.5232 | ND | 0.019775 | ND | 0.00098874 | ND | 0.25749 | ND | 0.00097761 | 0.2545 | -- | Not Applicable | --
| Sample 2 | [9 empty] | 2 |  | 337 | ND | 616.8448 | ND | 0.020012 | ND | 0.0010006 | ND | 0.26058 | ND |  |  | -- | Not Applicable | bar: Sample Speciation (ppmv) Run 1 to Detection Limits (ppmv) Run 4
| Background | ND | ND | ND | ND | 0 | 0 | 0 | 0 |  | 3 |  | 324 | ND | 593.0496 | ND | 0.01924 | ND | 0.00096202 | ND | 0.25036 | ND |  |  | -- | Not Applicable
[11 empty] | 4 |  | 323 | ND | 591.2192 | ND | 0.019181 | ND | 0.00095905 | ND | 0.24959 | ND |  |  | -- | Not Applicable
FTIR-Measured Compounds | [9 empty] | FTIR-Measured Compounds
Pollutant | Sample ID | Sample Speciation (ppmv) Run 1: Sample Speciation (ppmv); Detection Limits (ppmv) Run 1: Detection Limits (ppmv) | Pollutant | Run Number | Molecular Weight (g/gmol) | Sample Speciation with Detection Limits (SSDL) (ppmv)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Quantity of Pollutant Emitted (QPE) (lb)c |  | Emission Factor (lb/item)d |  | Emission Factor (lb/lb NEW)e |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)f |  | Minimum Detection Limit Based Concentration (mg/m3)
|  | Run 1 | Run 2 | Run 3 | Run 4 | Run 1 | Run 2 | Run 3 | Run 4 |  |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
NH3 | Sample 1 | 31.3 | 34.2 | 34.7 | 31.7 | -- | -- | -- | -- | NH3 | 1 | 17 | 31.3 | ND | 22.1354 | ND | 0.00071814 | ND | 3.5907e-05 | ND | 0.0093508 | ND | 3.7829e-05 | 0.0098479 | -- | Not Applicable | --
| Sample 2 | [9 empty] | 2 |  | 34.2 | ND | 24.1862 | ND | 0.00078468 | ND | 3.9234e-05 | ND | 0.010217 | ND |  |  | -- | Not Applicable | bar: Sample Speciation (ppmv) Run 1 to Detection Limits (ppmv) Run 4
| Background | ND | ND | ND | ND | 0.5 | 0.5 | 0.5 | 0.5 |  | 3 |  | 34.7 | ND | 24.5398 | ND | 0.00079615 | ND | 3.9808e-05 | ND | 0.01036 | ND |  |  | -- | Not Applicable
[11 empty] | 4 |  | 31.7 | ND | 22.4182 | ND | 0.00072732 | ND | 3.6366e-05 | ND | 0.0094639 | ND |  |  | -- | Not Applicable
a If the corresponding sample speciation value does not = "ND" or all active sample speciation values = "ND," Sample Speciation with Detection Limits (SSDL) = Sample Speciation. Otherwise, (i.e., if the corresponding sample speciation value = "ND" and
at least one of the active sample speciation values does not = "ND," Sample Speciation with Detection Limits (SSDL) = 0.5 * Detection Limit.
b Background Corrected Concentration (BCC) = Test Concentration - Background Concentration
c Quantity of Pollutant Emitted (QPE) =BCC * Initial Plume Volume
d Emission Factor = QPE / No. of Items
e Emission Factor = QPE / Net Explosive Weight
f If Average Emission Factor = "ND", Minimum Detection Limit Based Concentration = Test Minimum DL in mg/m3

## Acid Gases
Acid Gases Data and Emission Factor Calculations
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
Test Run Number | 1 | 2 | Background
Number of Items Tested | 20 | 20 | 0
Net Explosive Weight for All Items (lb) | 0.0768 | 0.0768 | NA
Initial Plume Volume (m3) | 14.716 | 14.716 | 14.716
Dilution Correction Factor | 0.89 | 0.888 | NA
Is Sample 1 Active? | Yes | Yes | Yes
Is Sample 2 Active? | No | No | No
Analytical Data Corrected for Volume at EPA Stationary Source Program Standard Conditions
Pollutant | Sample ID | Sample Speciation (mg/m3) |  | Detection Limits (mg/m3) |  | Pollutant | Run Number | Sample Speciation with Detection Limits (SSDL) (mg/m3)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC) (mg/m3)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3) | Test Concentration Relative Percent Difference
|  | Run 1 | Run 2 | Run 1 | Run 2 |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
Hydrogen Fluoride | Sample 1 | ND | ND | 0.18478 | 0.20524 | Hydrogen Fluoride | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.18478 | Not Applicable | 0.18478 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.20524 | Not Applicable
| Background | ND | ND | 0.20181 | 0.20181 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Hydrogen Chloride | Sample 1 | ND | ND | 0.17452 | 0.19497 | Hydrogen Chloride | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.17452 | Not Applicable | 0.17452 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.19497 | Not Applicable
| Background | ND | ND | 0.19172 | 0.19172 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Hydrogen Bromide | Sample 1 | ND | ND | 0.17452 | 0.19497 | Hydrogen Bromide | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.17452 | Not Applicable | 0.17452 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.19497 | Not Applicable
| Background | ND | ND | 0.19172 | 0.19172 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Nitric Acid | Sample 1 | 0.17452 | ND | ND | 0.19497 | Nitric Acid | 1 | 0.17452 | ND | 0.17452 | ND | 0.19609 | ND | 6.3617e-06 | ND | 3.1809e-07 | ND | 8.2835e-05 | ND | 2.4809e-07 | 6.4606e-05 | -- | Not Applicable | -- | 56.64%
| Sample 2 | ND | ND | ND | ND |  | 2 | 0.097487 | ND | 0.097487 | ND | 0.10978 | ND | 3.5617e-06 | ND | 1.7809e-07 | ND | 4.6376e-05 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.19172 | 0.19172 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Phosphoric Acid | Sample 1 | ND | ND | 0.17452 | 0.19497 | Phosphoric Acid | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.17452 | Not Applicable | 0.17452 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.19497 | Not Applicable
| Background | ND | ND | 0.19172 | 0.19172 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Sulfuric Acid | Sample 1 | ND | 0.26681 | 0.23611 | ND | Sulfuric Acid | 1 | 0.11806 | ND | 0.11806 | ND | 0.13265 | ND | 4.3035e-06 | ND | 2.1518e-07 | ND | 5.6035e-05 | ND | 3.5128e-07 | 9.148e-05 | -- | Not Applicable | -- | -77.3%
| Sample 2 | ND | ND | ND | ND |  | 2 | 0.26681 | ND | 0.26681 | ND | 0.30046 | ND | 9.7478e-06 | ND | 4.8739e-07 | ND | 0.00012693 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.19172 | 0.19172 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
a If the corresponding sample speciation value does not = "ND" or all active sample speciation values = "ND," Sample Speciation with Detection Limits (SSDL) = Sample Speciation. Otherwise, (i.e., if the corresponding sample speciation value = "ND" and
at least one of the active sample speciation values does not = "ND," Sample Speciation with Detection Limits (SSDL) = 0.5 * Detection Limit.
b Background Corrected Concentration (BCC) = Test Concentration - Background Concentration
c Dilution-Corrected Concentration (DCC) = BCC / Dilution Correction Factor
d Quantity of Pollutant Emitted (QPE) =BDCC * Initial Plume Volume
e Emission Factor = QPE / No. of Items
Test Run Number | Run 1 | Run 2 | Bkgd |  |  | f Emission Factor = QPE / Net Explosive Weight
Volume at Actual Conditions (m3) | 0.011846 | 0.010627 | 0.010598 |  |  | g If Average Emission Factor = "ND", Minimum Detection Limit Based Concentration = Test Minimum DL in mg/m3
Volume at EPA Stationary Source Program Standard Conditions (m3) | 0.01154 | 0.010356 | 0.010503
Volume Adjustment Factor | 1.0266 | 1.0262 | 1.0091
Original Analytical Data
Pollutant | Sample ID | Sample Speciation (mg/m3) |  | Detection Limits (mg/m3)
|  | Run 1 | Run 2 | Run 1 | Run 2
Hydrogen Fluoride | Sample 1 | ND | ND | 0.18 | 0.2
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.2 | 0.2
Hydrogen Chloride | Sample 1 | ND | ND | 0.17 | 0.19
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.19 | 0.19
Hydrogen Bromide | Sample 1 | ND | ND | 0.17 | 0.19
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.19 | 0.19
Nitric Acid | Sample 1 | 0.17 | ND | ND | 0.19
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.19 | 0.19
Phosphoric Acid | Sample 1 | ND | ND | 0.17 | 0.19
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.19 | 0.19
Sulfuric Acid | Sample 1 | ND | 0.26 | 0.23 | ND
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.19 | 0.19

## Cyanide
Cyanide Data and Emission Factor Calculations
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
Test Run Number | 1 | 2 | Background
Number of Items Tested | 20 | 20 | 0
Net Explosive Weight for All Items (lb) | 0.0768 | 0.0768 | NA
Initial Plume Volume (m3) | 14.716 | 14.716 | 14.716
Dilution Correction Factor | 0.89 | 0.888 | NA
Is Sample 1 Active? | Yes | Yes | Yes
Is Sample 2 Active? | No | No | No
Analytical Data Corrected for Volume at EPA Stationary Source Program Standard Conditions
Pollutant | Sample ID | Sample Speciation (mg/m3) |  | Detection Limits (mg/m3) |  | Pollutant | Run Number | Sample Speciation with Detection Limits (SSDL) (mg/m3)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC) (mg/m3)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3) | Test Concentration Relative Percent Difference
|  | Run 1 | Run 2 | Run 1 | Run 2 |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
Particulate Cyanide | Sample 1 | 0.89902 | 0.4587 | ND | ND | Particulate Cyanide | 1 | 0.89902 | ND | 0.89902 | ND | 1.0101 | ND | 3.2772e-05 | ND | 1.6386e-06 | ND | 0.00042672 | ND | 1.2383e-06 | 0.00032247 | -- | Not Applicable | -- | 64.86%
| Sample 2 | ND | ND | ND | ND |  | 2 | 0.4587 | ND | 0.4587 | ND | 0.51656 | ND | 1.6759e-05 | ND | 8.3794e-07 | ND | 0.00021821 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.012109 | 0.012109 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Hydrogen Cyanide | Sample 1 | 0.30708 | 2.9472 | ND | ND | Hydrogen Cyanide | 1 | 0.30708 | ND | 0.30708 | ND | 0.34503 | ND | 1.1194e-05 | ND | 5.597e-07 | ND | 0.00014575 | ND | 2.9718e-06 | 0.00077389 | -- | Not Applicable | -- | -162.3%
| Sample 2 | ND | ND | ND | ND |  | 2 | 2.9472 | ND | 2.9472 | ND | 3.3189 | ND | 0.00010768 | ND | 5.3838e-06 | ND | 0.001402 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.012109 | 0.012109 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
a If the corresponding sample speciation value does not = "ND" or all active sample speciation values = "ND," Sample Speciation with Detection Limits (SSDL) = Sample Speciation. Otherwise, (i.e., if the corresponding sample speciation value = "ND" and
at least one of the active sample speciation values does not = "ND," Sample Speciation with Detection Limits (SSDL) = 0.5 * Detection Limit.
b Background Corrected Concentration (BCC) = Test Concentration - Background Concentration
c Dilution-Corrected Concentration (DCC) = BCC / Dilution Correction Factor
d Quantity of Pollutant Emitted (QPE) =BDCC * Initial Plume Volume
e Emission Factor = QPE / No. of Items
Test Run Number | Run 1 | Run 2 | Bkgd |  |  | f Emission Factor = QPE / Net Explosive Weight
Volume at Actual Conditions (m3) | 0.020172 | 0.021509 | 0.021357 |  |  | g If Average Emission Factor = "ND", Minimum Detection Limit Based Concentration = Test Minimum DL in mg/m3
Volume at EPA Stationary Source Program Standard Conditions (m3) | 0.01997 | 0.02096 | 0.021165
Volume Adjustment Factor | 1.0101 | 1.0262 | 1.0091
Original Analytical Data
Pollutant | Sample ID | Sample Speciation (mg/m3) |  | Detection Limits (mg/m3)
|  | Run 1 | Run 2 | Run 1 | Run 2
Particulate Cyanide | Sample 1 | 0.89 | 0.447 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.012 | 0.012
Hydrogen Cyanide | Sample 1 | 0.304 | 2.872 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.012 | 0.012

## Particulates
Particulate Data and Emission Factor Calculations
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
Test Run Number | 1 | 2 | Background
Number of Items Tested | 20 | 20 | 0
Net Explosive Weight for All Items (lb) | 0.0768 | 0.0768 | NA
Initial Plume Volume (m3) | 14.716 | 14.716 | 14.716
Dilution Correction Factor | 0.89 | 0.888 | NA
Is Sample 1 Active? | Yes | Yes | No
Is Sample 2 Active? | No | No | No
Pollutant | Sample ID | Sample Speciation (mg/m3) |  | Detection Limits (mg/m3) |  | Pollutant | Run Number | Sample Speciation with Detection Limits (SSDL) (mg/m3)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC) (mg/m3)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3) | Test Concentration Relative Percent Difference
|  | Run 1 | Run 2 | Run 1 | Run 2 |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
TSP | Sample 1 | 38.6971 | 38.2001 | -- | -- | TSP | 1 | 38.6971 | ND | 38.6971 | ND | 43.4799 | ND | 0.0014106 | ND | 7.0531e-05 | ND | 0.018368 | ND | 7.0157e-05 | 0.01827 | -- | Not Applicable | -- | 1.293%
| Sample 2 | ND | ND | ND | ND |  | 2 | 38.2001 | ND | 38.2001 | ND | 43.0181 | ND | 0.0013956 | ND | 6.9782e-05 | ND | 0.018172 | ND |  |  | -- | Not Applicable
| Background | ND | ND | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
PM10 | Sample 1 | 37.0211 | 36.4668 | -- | -- | PM10 | 1 | 37.0211 | ND | 37.0211 | ND | 41.5968 | ND | 0.0013495 | ND | 6.7477e-05 | ND | 0.017572 | ND | 6.7046e-05 | 0.01746 | -- | Not Applicable | -- | 1.509%
| Sample 2 | ND | ND | ND | ND |  | 2 | 36.4668 | ND | 36.4668 | ND | 41.0662 | ND | 0.0013323 | ND | 6.6616e-05 | ND | 0.017348 | ND |  |  | -- | Not Applicable
| Background | ND | ND | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
PM2.5 | Sample 1 | 28.1069 | 28.0393 | -- | -- | PM2.5 | 1 | 28.1069 | ND | 28.1069 | ND | 31.5808 | ND | 0.0010246 | ND | 5.1229e-05 | ND | 0.013341 | ND | 5.1225e-05 | 0.01334 | -- | Not Applicable | -- | 0.2408%
| Sample 2 | ND | ND | ND | ND |  | 2 | 28.0393 | ND | 28.0393 | ND | 31.5758 | ND | 0.0010244 | ND | 5.1221e-05 | ND | 0.013339 | ND |  |  | -- | Not Applicable
| Background | ND | ND | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
a If the corresponding sample speciation value does not = "ND" or all active sample speciation values = "ND," Sample Speciation with Detection Limits (SSDL) = Sample Speciation. Otherwise, (i.e., if the corresponding sample speciation value = "ND" and
at least one of the active sample speciation values does not = "ND," Sample Speciation with Detection Limits (SSDL) = 0.5 * Detection Limit.
b Background Corrected Concentration (BCC) = Test Concentration - Background Concentration
c Dilution-Corrected Concentration (DCC) = BCC / Dilution Correction Factor
d Quantity of Pollutant Emitted (QPE) =BDCC * Initial Plume Volume
e Emission Factor = QPE / No. of Items
f Emission Factor = QPE / Net Explosive Weight
g If Average Emission Factor = "ND", Minimum Detection Limit Based Concentration = Test Minimum DL in mg/m3
Note on C18: TSP background conducted in order to obtain background for metals; however, no background conducted for PM10 or PM2.5. Background value of 0.121 was not used to determine TSP emission factor.
Note on D18: TSP background conducted in order to obtain background for metals; however, no background conducted for PM10 or PM2.5. Background value of 0.121 was not used to determine TSP emission factor.

## Carbon
Total Carbon Data and Emission Factor Calculations
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
Test Run Number | 1 | 2 | Background
Number of Items Tested | 20 | 20 | 0
Net Explosive Weight for All Items (lb) | 0.0768 | 0.0768 | NA
Initial Plume Volume (m3) | 14.716 | 14.716 | 14.716
Dilution Correction Factor | 0.89 | 0.888 | NA
Is Sample 1 Active? | Yes | Yes | Yes
Is Sample 2 Active? | No | No | No
Analytical Data Corrected for Volume at EPA Stationary Source Program Standard Conditions
Pollutant | Sample ID | Sample Speciation (mg/m3) |  | Detection Limits (mg/m3) |  | Pollutant | Run Number | Sample Speciation with Detection Limits (SSDL) (mg/m3)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC) (mg/m3)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3) | Test Concentration Relative Percent Difference
|  | Run 1 | Run 2 | Run 1 | Run 2 |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
Organic Carbon | Sample 1 | 2.4318 | 2.1172 | ND | ND | Organic Carbon | 1 | 2.4318 | ND | 2.3015 | ND | 2.586 | ND | 8.3898e-05 | ND | 4.1949e-06 | ND | 0.0010924 | ND | 3.9123e-06 | 0.0010188 | -- | Not Applicable | -- | 13.83%
Organic Carbon | Sample 2 | ND | ND | ND | ND | Organic Carbon | 2 | 2.1172 | ND | 1.987 | ND | 2.2376 | ND | 7.2595e-05 | ND | 3.6298e-06 | ND | 0.00094525 | ND |  |  | -- | Not Applicable
Organic Carbon | Background | 0.13022 | 0.13022 | ND | ND | Organic Carbon | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Elemental Carbon | Sample 1 | 0.54788 | 0.47115 | ND | ND | Elemental Carbon | 1 | 0.54788 | ND | 0.54788 | ND | 0.61559 | ND | 1.9972e-05 | ND | 9.9859e-07 | ND | 0.00026005 | ND | 9.2963e-07 | 0.00024209 | -- | Not Applicable | -- | 15.06%
Elemental Carbon | Sample 2 | ND | ND | ND | ND | Elemental Carbon | 2 | 0.47115 | ND | 0.47115 | ND | 0.53058 | ND | 1.7214e-05 | ND | 8.6068e-07 | ND | 0.00022413 | ND |  |  | -- | Not Applicable
Elemental Carbon | Background | ND | ND | 0.029263 | 0.029263 | Elemental Carbon | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
[6 empty] | Total Carbon | [13 empty] | 4.842e-06 | 0.0012609 |  |  | -- | --
[7 empty] | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
a If the corresponding sample speciation value does not = "ND" or all active sample speciation values = "ND," Sample Speciation with Detection Limits (SSDL) = Sample Speciation. Otherwise, (i.e., if the corresponding sample speciation value = "ND" and
at least one of the active sample speciation values does not = "ND," Sample Speciation with Detection Limits (SSDL) = 0.5 * Detection Limit.
b Background Corrected Concentration (BCC) = Test Concentration - Background Concentration
c Dilution-Corrected Concentration (DCC) = BCC / Dilution Correction Factor
d Quantity of Pollutant Emitted (QPE) =BDCC * Initial Plume Volume
Test Run Number | Run 1 | Run 2 | Bkgd |  |  | e Emission Factor = QPE / No. of Items
Volume at Actual Conditions (m3) | 0.04629 | 0.045978 | 0.04438 |  |  | f Emission Factor = QPE / Net Explosive Weight
Volume at EPA Stationary Source Program Standard Conditions (m3) | 0.045091 | 0.044805 | 0.043981 |  |  | g If Average Emission Factor = "ND", Minimum Detection Limit Based Concentration = Test Minimum DL in mg/m3
Volume Adjustment Factor | 1.0266 | 1.0262 | 1.0091
Original Analytical Data
Pollutant | Sample ID | Sample Speciation (mg/m3) |  | Detection Limits (mg/m3)
|  | Run 1 | Run 2 | Run 1 | Run 2
Organic Carbon | Sample 1 | 2.3688 | 2.0632 | ND | ND
Organic Carbon | Sample 2 | ND | ND | ND | ND
Organic Carbon | Background | 0.12905 | 0.12905 | ND | ND
Elemental Carbon | Sample 1 | 0.53369 | 0.45913 | ND | ND
Elemental Carbon | Sample 2 | ND | ND | ND | ND
Elemental Carbon | Background | ND | ND | 0.029 | 0.029

## Metals
Metals Data and Emission Factor Calculations
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
Test Run Number | 1 | 2 | Background
Number of Items Tested | 20 | 20 | 0
Net Explosive Weight for All Items (lb) | 0.0768 | 0.0768 | NA
Initial Plume Volume (m3) | 14.716 | 14.716 | 14.716
Dilution Correction Factor | 0.89 | 0.888 | NA
Is Sample 1 Active? | Yes | Yes | Yes
Is Sample 2 Active? | No | No | No
Metals Sample Volume (m3) Corrected to EPA Stationary Source Program Standard Conditions
Sample 1 | 0.44225 | 0.42695
Sample 2 | NA | NA
Background | 0.42534 | 0.42534
Pollutant | Sample ID | Sample Speciation (mg) |  | Detection Limits (mg) |  | Pollutant | Run Number | Sample Speciation with Detection Limits (SSDL) (mg)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC) (mg/m3)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3) | Test Concentration Relative Percent Difference
|  | Run 1 | Run 2 | Run 1 | Run 2 |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
Aluminum | Sample 1 | 0.0264 | 0.0261 | -- | -- | Aluminum | 1 | 0.0264 | ND | 0.059695 | ND | 0.067073 | ND | 2.1761e-06 | ND | 1.088e-07 | ND | 2.8334e-05 | ND | 1.1024e-07 | 2.8708e-05 | -- | Not Applicable | -- | -2.377%
| Sample 2 | ND | ND | ND | ND |  | 2 | 0.0261 | ND | 0.061132 | ND | 0.068842 | ND | 2.2335e-06 | ND | 1.1167e-07 | ND | 2.9081e-05 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.004 | 0.004 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Antimony | Sample 1 | 0.3228 | 0.3084 | -- | -- | Antimony | 1 | 0.3228 | ND | 0.72991 | ND | 0.82013 | ND | 2.6607e-05 | ND | 1.3304e-06 | ND | 0.00034645 | ND | 1.325e-06 | 0.00034504 | -- | Not Applicable | -- | 1.043%
| Sample 2 | ND | ND | ND | ND |  | 2 | 0.3084 | ND | 0.72234 | ND | 0.81344 | ND | 2.6391e-05 | ND | 1.3195e-06 | ND | 0.00034363 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.004 | 0.004 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Arsenic | Sample 1 | ND | ND | 0.004 | 0.004 | Arsenic | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0090448 | Not Applicable | 0.0090448 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0093688 | Not Applicable
| Background | ND | ND | 0.004 | 0.004 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Barium | Sample 1 | 0.1116 | 0.1105 | -- | -- | Barium | 1 | 0.1116 | ND | 0.25235 | ND | 0.28354 | ND | 9.1989e-06 | ND | 4.5994e-07 | ND | 0.00011978 | ND | 4.6637e-07 | 0.00012145 | -- | Not Applicable | -- | -2.53%
| Sample 2 | ND | ND | ND | ND |  | 2 | 0.1105 | ND | 0.25881 | ND | 0.29146 | ND | 9.4558e-06 | ND | 4.7279e-07 | ND | 0.00012312 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.002 | 0.002 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Beryllium | Sample 1 | ND | ND | 0.002 | 0.002 | Beryllium | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0045224 | Not Applicable | 0.0045224 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0046844 | Not Applicable
| Background | ND | ND | 0.002 | 0.002 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Cadmium | Sample 1 | ND | ND | 0.002 | 0.002 | Cadmium | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0045224 | Not Applicable | 0.0045224 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0046844 | Not Applicable
| Background | ND | ND | 0.002 | 0.002 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Calcium | Sample 1 | 0.0641 | 0.0569 | -- | -- | Calcium | 1 | 0.0641 | ND | 0.13201 | ND | 0.14833 | ND | 4.8122e-06 | ND | 2.4061e-07 | ND | 6.2659e-05 | ND | 2.3022e-07 | 5.9954e-05 | -- | Not Applicable | -- | 8.39%
| Sample 2 | ND | ND | ND | ND |  | 2 | 0.0569 | ND | 0.12034 | ND | 0.13552 | ND | 4.3967e-06 | ND | 2.1983e-07 | ND | 5.7248e-05 | ND |  |  | -- | Not Applicable
| Background | 0.0055 | 0.0055 | -- | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Chromium | Sample 1 | ND | ND | 0.001 | 0.001 | Chromium | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0022612 | Not Applicable | 0.0022612 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0023422 | Not Applicable
| Background | ND | ND | 0.001 | 0.001 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Cobalt | Sample 1 | ND | ND | 0.002 | 0.002 | Cobalt | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0045224 | Not Applicable | 0.0045224 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0046844 | Not Applicable
| Background | ND | ND | 0.002 | 0.002 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Copper | Sample 1 | 4.9591 | 4.9191 | -- | -- | Copper | 1 | 4.9591 | ND | 11.2135 | ND | 12.5994 | ND | 0.00040876 | ND | 2.0438e-05 | ND | 0.0053225 | ND | 2.0743e-05 | 0.0054017 | -- | Not Applicable | -- | -2.71%
| Sample 2 | ND | ND | ND | ND |  | 2 | 4.9191 | ND | 11.5216 | ND | 12.9747 | ND | 0.00042094 | ND | 2.1047e-05 | ND | 0.005481 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.004 | 0.004 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Lead | Sample 1 | 0.6792 | 0.6374 | -- | -- | Lead | 1 | 0.6792 | ND | 1.5358 | ND | 1.7256 | ND | 5.5985e-05 | ND | 2.7992e-06 | ND | 0.00072897 | ND | 2.7632e-06 | 0.00071959 | -- | Not Applicable | -- | 2.831%
| Sample 2 | ND | ND | ND | ND |  | 2 | 0.6374 | ND | 1.4929 | ND | 1.6812 | ND | 5.4544e-05 | ND | 2.7272e-06 | ND | 0.00071021 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.004 | 0.004 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Magnesium | Sample 1 | 0.1063 | 0.0937 | -- | -- | Magnesium | 1 | 0.1063 | ND | 0.24036 | ND | 0.27007 | ND | 8.762e-06 | ND | 4.381e-07 | ND | 0.00011409 | ND | 4.195e-07 | 0.00010925 | -- | Not Applicable | -- | 9.09%
| Sample 2 | ND | ND | ND | ND |  | 2 | 0.0937 | ND | 0.21946 | ND | 0.24715 | ND | 8.0182e-06 | ND | 4.0091e-07 | ND | 0.0001044 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.004 | 0.004 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Manganese | Sample 1 | ND | ND | 0.002 | 0.002 | Manganese | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0045224 | Not Applicable | 0.0045224 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0046844 | Not Applicable
| Background | ND | ND | 0.002 | 0.002 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Nickel | Sample 1 | ND | ND | 0.004 | 0.004 | Nickel | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0090448 | Not Applicable | 0.0090448 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0093688 | Not Applicable
| Background | ND | ND | 0.004 | 0.004 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Selenium | Sample 1 | ND | ND | 0.004 | 0.004 | Selenium | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0090448 | Not Applicable | 0.0090448 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0093688 | Not Applicable
| Background | ND | ND | 0.004 | 0.004 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Silver | Sample 1 | ND | ND | 0.001 | 0.001 | Silver | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0022612 | Not Applicable | 0.0022612 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0023422 | Not Applicable
| Background | ND | ND | 0.001 | 0.001 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Thallium | Sample 1 | ND | 0.0045 | 0.004 | -- | Thallium | 1 | 0.002 | ND | 0.0045224 | ND | 0.0050813 | ND | 1.6485e-07 | ND | 8.2427e-09 | ND | 2.1465e-06 | ND | 1.3748e-08 | 3.5803e-06 | -- | Not Applicable | -- | -79.9%
| Sample 2 | ND | ND | ND | ND |  | 2 | 0.0045 | ND | 0.01054 | ND | 0.011869 | ND | 3.8508e-07 | ND | 1.9254e-08 | ND | 5.014e-06 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.004 | 0.004 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Vanadium | Sample 1 | ND | ND | 0.002 | 0.002 | Vanadium | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0045224 | Not Applicable | 0.0045224 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0046844 | Not Applicable
| Background | ND | ND | 0.002 | 0.002 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Zinc | Sample 1 | 0.5632 | 0.5704 | -- | -- | Zinc | 1 | 0.5632 | ND | 1.2735 | ND | 1.4309 | ND | 4.6423e-05 | ND | 2.3211e-06 | ND | 0.00060447 | ND | 2.3808e-06 | 0.00062001 | -- | Not Applicable | -- | -4.79%
| Sample 2 | ND | ND | ND | ND |  | 2 | 0.5704 | ND | 1.336 | ND | 1.5045 | ND | 4.8811e-05 | ND | 2.4405e-06 | ND | 0.00063556 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.002 | 0.002 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
a If the corresponding sample speciation value does not = "ND" or all active sample speciation values = "ND," Sample Speciation with Detection Limits (SSDL) = Sample Speciation. Otherwise, (i.e., if the corresponding sample speciation value = "ND" and
at least one of the active sample speciation values does not = "ND," Sample Speciation with Detection Limits (SSDL) = 0.5 * Detection Limit.
b Background Corrected Concentration (BCC) = Test Mass/Test Sample Volume - Background Mass/ Background Sample Volume
c Dilution-Corrected Concentration (DCC) = BCC / Dilution Correction Factor
d Quantity of Pollutant Emitted (QPE) =BDCC * Initial Plume Volume
e Emission Factor = QPE / No. of Items
f Emission Factor = QPE / Net Explosive Weight
g If Average Emission Factor = "ND", Minimum Detection Limit Based Concentration = Test DL/Test Sample volume in mg/m3

## Hexavalent Chromium
Hexavalent Chromium Data and Emission Factor Calculations
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
Test Run Number | 1 | 2 | Background
Number of Items Tested | 20 | 20 | 0
Net Explosive Weight for All Items (lb) | 0.0768 | 0.0768 | NA
Initial Plume Volume (m3) | 14.716 | 14.716 | 14.716
Dilution Correction Factor | 0.89 | 0.888 | NA
Is Sample 1 Active? | Yes | Yes | Yes
Is Sample 2 Active? | No | No | No
Analytical Data Corrected for Volume at EPA Stationary Source Program Standard Conditions
Pollutant | Sample ID | Sample Speciation (mg/m3) |  | Detection Limits (mg/m3) |  | Pollutant | Run Number | Sample Speciation with Detection Limits (SSDL) (mg/m3)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC) (mg/m3)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3) | Test Concentration Relative Percent Difference
|  | Run 1 | Run 2 | Run 1 | Run 2 |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
Hexavalent Chromium | Sample 1 | ND | ND | 0.00061697 | 0.00060545 | Hexavalent Chromium | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00061697 | Not Applicable | 0.00060545 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00060545 | Not Applicable
| Background | 0.00058425 | 0.00058425 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
a If the corresponding sample speciation value does not = "ND" or all active sample speciation values = "ND," Sample Speciation with Detection Limits (SSDL) = Sample Speciation. Otherwise, (i.e., if the corresponding sample speciation value = "ND" and
at least one of the active sample speciation values does not = "ND," Sample Speciation with Detection Limits (SSDL) = 0.5 * Detection Limit.
b Background Corrected Concentration (BCC) = Test Concentration - Background Concentration
c Dilution-Corrected Concentration (DCC) = BCC / Dilution Correction Factor
d Quantity of Pollutant Emitted (QPE) =BDCC * Initial Plume Volume
e Emission Factor = QPE / No. of Items
f Emission Factor = QPE / Net Explosive Weight
g If Average Emission Factor = "ND", Minimum Detection Limit Based Concentration = Test Minimum DL in mg/m3
Test Run Number | Run 1 | Run 2 | Bkgd
Volume at Actual Conditions (m3) | 0.066567 | 0.067821 | 0.070767
Volume at EPA Stationary Source Program Standard Conditions (m3) | 0.064843 | 0.066091 | 0.070132
Volume Adjustment Factor | 1.0266 | 1.0262 | 1.0091
Original Analytical Data
Pollutant | Sample ID | Sample Speciation (mg/m3) |  | Detection Limits (mg/m3)
|  | Run 1 | Run 2 | Run 1 | Run 2
Hexavalent Chromium | Sample 1 | ND | ND | 0.000601 | 0.00059
| Sample 2 | ND | ND | ND | ND
| Background | 0.000579 | 0.000579 | ND | ND

## Aldehyde
TO-11 Carbonyl Data and Emission Factor Calculations
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
Test Run Number | 1 | 2 | Background
Number of Items Tested | 20 | 20 | 0
Net Explosive Weight for All Items (lb) | 0.0768 | 0.0768 | NA
Initial Plume Volume (m3) | 14.716 | 14.716 | 14.716
Dilution Correction Factor | 0.89 | 0.888 | NA
Is Sample 1 Active? | Yes | Yes | Yes
Is Sample 2 Active? | No | No | No
Analytical Data Corrected for Volume at EPA Stationary Source Program Standard Conditions
Pollutant | Sample ID | Sample Speciation (mg/m3) |  | Detection Limits (mg/m3) |  | Pollutant | Run Number | Sample Speciation with Detection Limits (SSDL) (mg/m3)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC) (mg/m3)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3) | Test Concentration Relative Percent Difference
|  | Run 1 | Run 2 | Run 1 | Run 2 |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
Formaldehyde | Sample 1 | 0.10266 | 0.10159 | ND | ND | Formaldehyde | 1 | 0.10266 | ND | 0.085504 | ND | 0.096072 | ND | 3.1169e-06 | ND | 1.5584e-07 | ND | 4.0584e-05 | ND | 1.5505e-07 | 4.0376e-05 | -- | Not Applicable | -- | 1.044%
| Sample 2 | ND | ND | ND | ND |  | 2 | 0.10159 | ND | 0.084438 | ND | 0.095088 | ND | 3.0849e-06 | ND | 1.5425e-07 | ND | 4.0169e-05 | ND |  |  | -- | Not Applicable
| Background | 0.017154 | 0.017154 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Acetaldehyde | Sample 1 | 0.067754 | 0.093383 | ND | ND | Acetaldehyde | 1 | 0.067754 | ND | 0.054636 | ND | 0.061389 | ND | 1.9917e-06 | ND | 9.9583e-08 | ND | 2.5933e-05 | ND | 1.231e-07 | 3.2058e-05 | -- | Not Applicable | -- | -31.81%
| Sample 2 | ND | ND | ND | ND |  | 2 | 0.093383 | ND | 0.080265 | ND | 0.090388 | ND | 2.9325e-06 | ND | 1.4662e-07 | ND | 3.8183e-05 | ND |  |  | -- | Not Applicable
| Background | 0.013118 | 0.013118 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Acetone | Sample 1 | 0.12319 | 0.038995 | ND | ND | Acetone | 1 | 0.12319 | ND | 0.077782 | ND | 0.087395 | ND | 2.8354e-06 | ND | 1.4177e-07 | ND | 3.6919e-05 | ND | 7.0884e-08 | 1.8459e-05 | -- | Not Applicable | -- | 103.8%
| Sample 2 | ND | ND | ND | ND |  | 2 | 0.038995 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 0.045408 | 0.045408 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Acrolein | Sample 1 | ND | ND | 0.020532 | 0.01334 | Acrolein | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.020532 | Not Applicable | 0.01334 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01334 | Not Applicable
| Background | ND | ND | 0.0074671 | 0.0074671 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Proprionaldehyde | Sample 1 | ND | ND | 0.020532 | 0.020524 | Proprionaldehyde | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.020532 | Not Applicable | 0.020524 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.020524 | Not Applicable
| Background | ND | ND | 0.0074671 | 0.0074671 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Crotonaldehyde | Sample 1 | ND | ND | 0.020532 | 0.020524 | Crotonaldehyde | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.020532 | Not Applicable | 0.020524 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.020524 | Not Applicable
| Background | ND | ND | 0.0074671 | 0.0074671 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Butyraldehyde | Sample 1 | ND | ND | 0.020532 | 0.020524 | Butyraldehyde | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.020532 | Not Applicable | 0.020524 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.020524 | Not Applicable
| Background | ND | ND | 0.0074671 | 0.0074671 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzaldehyde | Sample 1 | ND | ND | 0.020532 | 0.020524 | Benzaldehyde | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.020532 | Not Applicable | 0.020524 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.020524 | Not Applicable
| Background | ND | ND | 0.0074671 | 0.0074671 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Isovaleraldehyde | Sample 1 | ND | ND | 0.020532 | 0.020524 | Isovaleraldehyde | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.020532 | Not Applicable | 0.020524 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.020524 | Not Applicable
| Background | ND | ND | 0.0074671 | 0.0074671 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Valeraldehyde | Sample 1 | ND | ND | 0.020532 | 0.020524 | Valeraldehyde | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.020532 | Not Applicable | 0.020524 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.020524 | Not Applicable
| Background | ND | ND | 0.0074671 | 0.0074671 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
o,m,p-Tolualdehyde | Sample 1 | ND | ND | 0.061595 | 0.061571 | o,m,p-Tolualdehyde | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.061595 | Not Applicable | 0.061571 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.061571 | Not Applicable
| Background | ND | ND | 0.022199 | 0.022199 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Hexaldehyde | Sample 1 | ND | ND | 0.020532 | 0.020524 | Hexaldehyde | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.020532 | Not Applicable | 0.020524 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.020524 | Not Applicable
| Background | ND | ND | 0.0074671 | 0.0074671 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,5-Dimethylbenzaldehyde | Sample 1 | ND | ND | 0.020532 | 0.020524 | 2,5-Dimethylbenzaldehyde | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.020532 | Not Applicable | 0.020524 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.020524 | Not Applicable
| Background | ND | ND | 0.0074671 | 0.0074671 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
a If the corresponding sample speciation value does not = "ND" or all active sample speciation values = "ND," Sample Speciation with Detection Limits (SSDL) = Sample Speciation. Otherwise, (i.e., if the corresponding sample speciation value = "ND" and
at least one of the active sample speciation values does not = "ND," Sample Speciation with Detection Limits (SSDL) = 0.5 * Detection Limit.
b Background Corrected Concentration (BCC) = Test Concentration - Background Concentration
c Dilution-Corrected Concentration (DCC) = BCC / Dilution Correction Factor
d Quantity of Pollutant Emitted (QPE) =BDCC * Initial Plume Volume
e Emission Factor = QPE / No. of Items
Test Run Number | Run 1 | Run 2 | Bkgd |  |  | f Emission Factor = QPE / Net Explosive Weight
Volume at Actual Conditions (m3) | 0.014949 | 0.023987 | 0.04048 |  |  | g If Average Emission Factor = "ND", Minimum Detection Limit Based Concentration = Test Minimum DL in mg/m3
Volume at EPA Stationary Source Program Standard Conditions (m3) | 0.014562 | 0.023375 | 0.040117
Volume Adjustment Factor | 1.0266 | 1.0262 | 1.0091
Original Analytical Data
Pollutant | Sample ID | Sample Speciation (mg/m3) |  | Detection Limits (mg/m3)
|  | Run 1 | Run 2 | Run 1 | Run 2
Formaldehyde | Sample 1 | 0.1 | 0.099 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 0.017 | 0.017 | ND | ND
Acetaldehyde | Sample 1 | 0.066 | 0.091 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 0.013 | 0.013 | ND | ND
Acetone | Sample 1 | 0.12 | 0.038 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 0.045 | 0.045 | ND | ND
Acrolein | Sample 1 | ND | ND | 0.02 | 0.013
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.0074 | 0.0074
Proprionaldehyde | Sample 1 | ND | ND | 0.02 | 0.02
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.0074 | 0.0074
Crotonaldehyde | Sample 1 | ND | ND | 0.02 | 0.02
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.0074 | 0.0074
Butyraldehyde | Sample 1 | ND | ND | 0.02 | 0.02
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.0074 | 0.0074
Benzaldehyde | Sample 1 | ND | ND | 0.02 | 0.02
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.0074 | 0.0074
Isovaleraldehyde | Sample 1 | ND | ND | 0.02 | 0.02
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.0074 | 0.0074
Valeraldehyde | Sample 1 | ND | ND | 0.02 | 0.02
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.0074 | 0.0074
o,m,p-Tolualdehyde | Sample 1 | ND | ND | 0.06 | 0.06
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.022 | 0.022
Hexaldehyde | Sample 1 | ND | ND | 0.02 | 0.02
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.0074 | 0.0074
2,5-Dimethylbenzaldehyde | Sample 1 | ND | ND | 0.02 | 0.02
| Sample 2 | ND | ND | ND | ND
| Background | ND | ND | 0.0074 | 0.0074

## VOC &VOC TIC
VOC & VOC TIC Data and Emission Factor Calculations
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
Test Run Number | 1 | 2 | Background
Number of Items Tested | 20 | 20 | 0
Net Explosive Weight for All Items (lb) | 0.0768 | 0.0768 | NA
Initial Plume Volume (m3) | 14.716 | 14.716 | 14.716
Dilution Correction Factor | 0.89 | 0.888 | NA
Is Sample 1 Active? | Yes | Yes | Yes
Is Sample 2 Active? | No | No | No
Pollutant | Sample ID | Sample Speciation |  | Detection Limits |  | Pollutant | Run Number | Molecular Weight (g/gmol) | Sample Speciation with Detection Limits (SSDL) (ppbv)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3) | Test Concentration Relative Percent Difference
|  | Run 1 | Run 2 | Run 1 | Run 2
|  | (ppbv) | (ppbv) | (ppbv) | (ppbv) |  |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
Propene | Sample 1 | 20 | 16 | -- | -- | Propene | 1 | 42 | 20 | ND | 0.034944 | ND | 0.039263 | ND | 1.2738e-06 | ND | 6.3691e-08 | ND | 1.6586e-05 | ND | 5.7379e-08 | 1.4942e-05 | -- | Not Applicable | -- | 22.22%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 16 | ND | 0.027955 | ND | 0.031481 | ND | 1.0213e-06 | ND | 5.1067e-08 | ND | 1.3299e-05 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.09 | 0.09 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Dichlorodifluoromethane | Sample 1 | 0.6 | 0.6 | -- | -- | Dichlorodifluoromethane | 1 | 120 | 0.6 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | 0 | -- | Not Applicable | -- | 0%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 0.6 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 0.6 | 0.6 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Chlorodifluoromethane | Sample 1 | ND | ND | 0.07 | 0.07 | Chlorodifluoromethane | 1 | 86.47 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0002518 | Not Applicable | 0.0002518 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0002518 | Not Applicable
| Background | ND | ND | 0.07 | 0.07 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Freon 114 | Sample 1 | ND | ND | 0.07 | 0.07 | Freon 114 | 1 | 170.93 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00049775 | Not Applicable | 0.00049775 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00049775 | Not Applicable
| Background | ND | ND | 0.07 | 0.07 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Chloromethane | Sample 1 | 0.6 | 0.6 | -- | -- | Chloromethane | 1 | 50.49 | 0.6 | ND | 0.0012602 | ND | 0.001416 | ND | 4.5939e-08 | ND | 2.297e-09 | ND | 5.9817e-07 | ND | 2.2995e-09 | 5.9884e-07 | -- | Not Applicable | -- | 0%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 0.6 | ND | 0.0012602 | ND | 0.0014192 | ND | 4.6043e-08 | ND | 2.3021e-09 | ND | 5.9951e-07 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.14 | 0.14 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Vinyl Chloride | Sample 1 | ND | ND | 0.1 | 0.1 | Vinyl Chloride | 1 | 62.5 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00026 | Not Applicable | 0.00026 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00026 | Not Applicable
| Background | ND | ND | 0.1 | 0.1 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,3-Butadiene | Sample 1 | 3 | 2 | -- | -- | 1,3-Butadiene | 1 | 54 | 3 | ND | 0.0067392 | ND | 0.0075721 | ND | 2.4566e-07 | ND | 1.2283e-08 | ND | 3.1988e-06 | ND | 1.0245e-08 | 2.668e-06 | -- | Not Applicable | -- | 40%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 2 | ND | 0.0044928 | ND | 0.0050595 | ND | 1.6414e-07 | ND | 8.2072e-09 | ND | 2.1373e-06 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.11 | 0.11 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Bromomethane | Sample 1 | ND | ND | 0.09 | 0.09 | Bromomethane | 1 | 94.94 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00035546 | Not Applicable | 0.00035546 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00035546 | Not Applicable
| Background | ND | ND | 0.09 | 0.09 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Chloroethane | Sample 1 | ND | ND | 0.23 | 0.23 | Chloroethane | 1 | 64.52 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00061733 | Not Applicable | 0.00061733 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00061733 | Not Applicable
| Background | ND | ND | 0.23 | 0.23 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Dichlorofluoromethane | Sample 1 | ND | ND | 0.05 | 0.05 | Dichlorofluoromethane | 1 | 102.92 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00021407 | Not Applicable | 0.00021407 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00021407 | Not Applicable
| Background | ND | ND | 0.05 | 0.05 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Trichloroflouromethane | Sample 1 | 0.3 | 0.3 | -- | -- | Trichloroflouromethane | 1 | 137.37 | 0.3 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | 0 | -- | Not Applicable | -- | 0%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 0.3 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 0.3 | 0.3 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Pentane | Sample 1 | 0.5 | ND | -- | 0.07 | Pentane | 1 | 72.15 | 0.5 | ND | 0.00090043 | ND | 0.0010117 | ND | 3.2823e-08 | ND | 1.6412e-09 | ND | 4.2739e-07 | ND | 8.2059e-10 | 2.1369e-07 | -- | Not Applicable | -- | 173.8%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 0.035 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 0.2 | 0.2 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Acrolein | Sample 1 | 1 | ND | -- | 0.49 | Acrolein | 1 | 56.06 | 1 | ND | 0.0023321 | ND | 0.0026203 | ND | 8.5012e-08 | ND | 4.2506e-09 | ND | 1.1069e-06 | ND | 2.6472e-09 | 6.8937e-07 | -- | Not Applicable | -- | 121.3%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 0.245 | ND | 0.00057136 | ND | 0.00064343 | ND | 2.0875e-08 | ND | 1.0437e-09 | ND | 2.7181e-07 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.49 | 0.49 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,1-Dichlorethene | Sample 1 | ND | ND | 0.05 | 0.05 | 1,1-Dichlorethene | 1 | 98.96 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00020584 | Not Applicable | 0.00020584 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00020584 | Not Applicable
| Background | ND | ND | 0.05 | 0.05 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Freon 113 | Sample 1 | ND | ND | 0.07 | 0.07 | Freon 113 | 1 | 187.83 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00054696 | Not Applicable | 0.00054696 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00054696 | Not Applicable
| Background | ND | ND | 0.07 | 0.07 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Acetone | Sample 1 | 9 | 9 | -- | -- | Acetone | 1 | 58.08 | 9 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | 0 | -- | Not Applicable | -- | 0%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 9 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 9 | 9 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Methyl Iodide | Sample 1 | ND | ND | 0.06 | 0.06 | Methyl Iodide | 1 | 141.94 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00035428 | Not Applicable | 0.00035428 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00035428 | Not Applicable
| Background | ND | ND | 0.06 | 0.06 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Carbon Disulfide | Sample 1 | ND | ND | 0.1 | 0.1 | Carbon Disulfide | 1 | 76 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00031616 | Not Applicable | 0.00031616 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00031616 | Not Applicable
| Background | ND | ND | 0.1 | 0.1 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Acetonitrile | Sample 1 | 79 | 72 | -- | -- | Acetonitrile | 1 | 41 | 79 | ND | 0.13474 | ND | 0.1514 | ND | 4.9118e-06 | ND | 2.4559e-07 | ND | 6.3955e-05 | ND | 2.3496e-07 | 6.1187e-05 | -- | Not Applicable | -- | 9.272%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 72 | ND | 0.1228 | ND | 0.13829 | ND | 4.4866e-06 | ND | 2.2433e-07 | ND | 5.842e-05 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.17 | 0.17 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
3-Chloropropene | Sample 1 | ND | ND | 0.28 | 0.28 | 3-Chloropropene | 1 | 76.53 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00089142 | Not Applicable | 0.00089142 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00089142 | Not Applicable
| Background | ND | ND | 0.28 | 0.28 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Methylene chloride | Sample 1 | 13 | 14 | -- | -- | Methylene chloride | 1 | 85 | 13 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | 0 | -- | Not Applicable | -- | -7.407%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 14 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 24 | 24 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
tert-Butyl Alcohol | Sample 1 | ND | ND | 0.13 | 0.13 | tert-Butyl Alcohol | 1 | 74.12 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00040084 | Not Applicable | 0.00040084 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00040084 | Not Applicable
| Background | ND | ND | 0.13 | 0.13 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Acrylonitrile | Sample 1 | 5 | 5 | -- | -- | Acrylonitrile | 1 | 53 | 5 | ND | 0.011024 | ND | 0.012387 | ND | 4.0186e-07 | ND | 2.0093e-08 | ND | 5.2325e-06 | ND | 2.0116e-08 | 5.2384e-06 | -- | Not Applicable | -- | 0%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 5 | ND | 0.011024 | ND | 0.012414 | ND | 4.0276e-07 | ND | 2.0138e-08 | ND | 5.2443e-06 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.27 | 0.27 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
trans-1,2,-Dichloroethene | Sample 1 | ND | ND | 0.12 | 0.12 | trans-1,2,-Dichloroethene | 1 | 97 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00048422 | Not Applicable | 0.00048422 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00048422 | Not Applicable
| Background | ND | ND | 0.12 | 0.12 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Methyl t-Butyl Ether | Sample 1 | 0.5 | ND | -- | 0.1 | Methyl t-Butyl Ether | 1 | 88.15 | 0.5 | ND | 0.00073341 | ND | 0.00082405 | ND | 2.6735e-08 | ND | 1.3367e-09 | ND | 3.4811e-07 | ND | 6.6837e-10 | 1.7406e-07 | -- | Not Applicable | -- | 163.6%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 0.05 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 0.3 | 0.3 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Hexane | Sample 1 | 3 | 2 | -- | -- | Hexane | 1 | 86.18 | 3 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | 0 | -- | Not Applicable | -- | 40%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 2 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 5 | 5 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,1-Dichloroethane | Sample 1 | ND | ND | 0.09 | 0.09 | 1,1-Dichloroethane | 1 | 99 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00037066 | Not Applicable | 0.00037066 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00037066 | Not Applicable
| Background | ND | ND | 0.09 | 0.09 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Vinyl Acetate | Sample 1 | ND | ND | 0.11 | 0.11 | Vinyl Acetate | 1 | 86.09 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00039395 | Not Applicable | 0.00039395 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00039395 | Not Applicable
| Background | ND | ND | 0.11 | 0.11 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
cis-1,2-Dichloroethene | Sample 1 | ND | ND | 0.05 | 0.05 | cis-1,2-Dichloroethene | 1 | 96.94 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00020164 | Not Applicable | 0.00020164 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00020164 | Not Applicable
| Background | ND | ND | 0.05 | 0.05 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2-Butanone | Sample 1 | 0.9 | 0.5 | -- | -- | 2-Butanone | 1 | 72 | 0.9 | ND | 0.0026957 | ND | 0.0030289 | ND | 9.8266e-08 | ND | 4.9133e-09 | ND | 1.2795e-06 | ND | 3.8245e-09 | 9.9597e-07 | -- | Not Applicable | -- | 57.14%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 0.5 | ND | 0.0014976 | ND | 0.0016865 | ND | 5.4715e-08 | ND | 2.7357e-09 | ND | 7.1243e-07 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.27 | 0.27 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Ethyl Acetate | Sample 1 | ND | ND | 0.37 | 0.37 | Ethyl Acetate | 1 | 88.11 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0013562 | Not Applicable | 0.0013562 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0013562 | Not Applicable
| Background | ND | ND | 0.37 | 0.37 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Methyl Acrylate | Sample 1 | ND | ND | 0.07 | 0.07 | Methyl Acrylate | 1 | 86.1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00025072 | Not Applicable | 0.00025072 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00025072 | Not Applicable
| Background | ND | ND | 0.07 | 0.07 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Chloroform | Sample 1 | ND | ND | 0.08 | 0.08 | Chloroform | 1 | 119 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00039603 | Not Applicable | 0.00039603 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00039603 | Not Applicable
| Background | ND | ND | 0.08 | 0.08 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,1,1-Trichloroethane | Sample 1 | ND | ND | 0.08 | 0.08 | 1,1,1-Trichloroethane | 1 | 133.41 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00044399 | Not Applicable | 0.00044399 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00044399 | Not Applicable
| Background | ND | ND | 0.08 | 0.08 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Carbontetrachloride | Sample 1 | ND | ND | 0.05 | 0.05 | Carbontetrachloride | 1 | 154 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00032032 | Not Applicable | 0.00032032 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00032032 | Not Applicable
| Background | ND | ND | 0.05 | 0.05 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,2-Dichloroethane | Sample 1 | 0.5 | 0.5 | -- | -- | 1,2-Dichloroethane | 1 | 99 | 0.5 | ND | 0.0020592 | ND | 0.0023137 | ND | 7.5064e-08 | ND | 3.7532e-09 | ND | 9.774e-07 | ND | 3.7574e-09 | 9.785e-07 | -- | Not Applicable | -- | 0%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 0.5 | ND | 0.0020592 | ND | 0.0023189 | ND | 7.5233e-08 | ND | 3.7617e-09 | ND | 9.796e-07 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.05 | 0.05 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzene | Sample 1 | 33 | 32 | -- | -- | Benzene | 1 | 78 | 33 | ND | 0.10708 | ND | 0.12031 | ND | 3.9033e-06 | ND | 1.9517e-07 | ND | 5.0825e-05 | ND | 1.9242e-07 | 5.011e-05 | -- | Not Applicable | -- | 3.077%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 32 | ND | 0.10383 | ND | 0.11693 | ND | 3.7936e-06 | ND | 1.8968e-07 | ND | 4.9395e-05 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.07 | 0.07 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Isooctane | Sample 1 | ND | ND | 0.04 | 0.04 | Isooctane | 1 | 114.23 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00019008 | Not Applicable | 0.00019008 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00019008 | Not Applicable
| Background | ND | ND | 0.04 | 0.04 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Heptane | Sample 1 | ND | ND | 0.05 | 0.05 | Heptane | 1 | 100.2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00020842 | Not Applicable | 0.00020842 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00020842 | Not Applicable
| Background | ND | ND | 0.05 | 0.05 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Trichloroethane | Sample 1 | ND | ND | 0.06 | 0.06 | Trichloroethane | 1 | 119.38 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00029797 | Not Applicable | 0.00029797 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00029797 | Not Applicable
| Background | ND | ND | 0.06 | 0.06 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Ethyl Acrylate | Sample 1 | ND | ND | 0.04 | 0.04 | Ethyl Acrylate | 1 | 100.12 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0001666 | Not Applicable | 0.0001666 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0001666 | Not Applicable
| Background | ND | ND | 0.04 | 0.04 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,2-Dichloropropane | Sample 1 | ND | ND | 0.07 | 0.07 | 1,2-Dichloropropane | 1 | 113 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00032906 | Not Applicable | 0.00032906 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00032906 | Not Applicable
| Background | ND | ND | 0.07 | 0.07 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Methyl Methacrylate | Sample 1 | ND | ND | 0.11 | 0.11 | Methyl Methacrylate | 1 | 100.12 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00045815 | Not Applicable | 0.00045815 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00045815 | Not Applicable
| Background | ND | ND | 0.11 | 0.11 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Dibromomethane | Sample 1 | ND | ND | 0.07 | 0.07 | Dibromomethane | 1 | 173.83 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00050619 | Not Applicable | 0.00050619 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00050619 | Not Applicable
| Background | ND | ND | 0.07 | 0.07 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,4-Dioxane | Sample 1 | ND | ND | 0.09 | 0.09 | 1,4-Dioxane | 1 | 88.11 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00032988 | Not Applicable | 0.00032988 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00032988 | Not Applicable
| Background | ND | ND | 0.09 | 0.09 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Bromodichloromethane | Sample 1 | ND | ND | 0.06 | 0.06 | Bromodichloromethane | 1 | 163.83 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00040892 | Not Applicable | 0.00040892 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00040892 | Not Applicable
| Background | ND | ND | 0.06 | 0.06 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
cis-1,3-Dichloropropene | Sample 1 | ND | ND | 0.06 | 0.06 | cis-1,3-Dichloropropene | 1 | 110.9706 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00027698 | Not Applicable | 0.00027698 | ND
cis-1,3-Dichloropropene | Sample 2 | ND | ND | ND | ND | cis-1,3-Dichloropropene | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00027698 | Not Applicable
cis-1,3-Dichloropropene | Background | ND | ND | 0.06 | 0.06 | cis-1,3-Dichloropropene | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
4-Methyl-2-pentanone | Sample 1 | ND | ND | 0.08 | 0.08 | 4-Methyl-2-pentanone | 1 | 100 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0003328 | Not Applicable | 0.0003328 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0003328 | Not Applicable
| Background | ND | ND | 0.08 | 0.08 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Toluene | Sample 1 | 3 | 2 | -- | -- | Toluene | 1 | 92 | 3 | ND | 0.010333 | ND | 0.011611 | ND | 3.7669e-07 | ND | 1.8834e-08 | ND | 4.9048e-06 | ND | 1.536e-08 | 3.9999e-06 | -- | Not Applicable | -- | 40%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 2 | ND | 0.0065062 | ND | 0.0073268 | ND | 2.3771e-07 | ND | 1.1885e-08 | ND | 3.0951e-06 | ND |  |  | -- | Not Applicable
| Background | 0.3 | 0.3 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Octane | Sample 1 | ND | ND | 0.04 | 0.04 | Octane | 1 | 114.23 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00019008 | Not Applicable | 0.00019008 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00019008 | Not Applicable
| Background | ND | ND | 0.04 | 0.04 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
trans-1,3-Dichloropropene | Sample 1 | ND | ND | 0.05 | 0.05 | trans-1,3-Dichloropropene | 1 | 110.97 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00023082 | Not Applicable | 0.00023082 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00023082 | Not Applicable
| Background | ND | ND | 0.05 | 0.05 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Ethyl Methacrylate | Sample 1 | ND | ND | 0.07 | 0.07 | Ethyl Methacrylate | 1 | 114.14 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00033238 | Not Applicable | 0.00033238 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00033238 | Not Applicable
| Background | ND | ND | 0.07 | 0.07 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,1,2-Trichloroethane | Sample 1 | ND | ND | 0.05 | 0.05 | 1,1,2-Trichloroethane | 1 | 133 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00027664 | Not Applicable | 0.00027664 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00027664 | Not Applicable
| Background | ND | ND | 0.05 | 0.05 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Tertrachloroethene | Sample 1 | ND | ND | 0.06 | 0.06 | Tertrachloroethene | 1 | 165.83 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00041391 | Not Applicable | 0.00041391 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00041391 | Not Applicable
| Background | ND | ND | 0.06 | 0.06 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2-Hexanone | Sample 1 | ND | ND | 0.07 | 0.07 | 2-Hexanone | 1 | 100 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0002912 | Not Applicable | 0.0002912 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0002912 | Not Applicable
| Background | ND | ND | 0.07 | 0.07 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Dibromochloromethane | Sample 1 | ND | ND | 0.06 | 0.06 | Dibromochloromethane | 1 | 208.21 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00051969 | Not Applicable | 0.00051969 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00051969 | Not Applicable
| Background | ND | ND | 0.06 | 0.06 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,2-Dibromoethane | Sample 1 | ND | ND | 0.05 | 0.05 | 1,2-Dibromoethane | 1 | 188 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00039104 | Not Applicable | 0.00039104 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00039104 | Not Applicable
| Background | ND | ND | 0.05 | 0.05 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Chlorobenzene | Sample 1 | ND | 0.2 | 0.05 | -- | Chlorobenzene | 1 | 113 | 0.025 | ND | 0.00011752 | ND | 0.00013204 | ND | 4.284e-09 | ND | 2.142e-10 | ND | 5.5781e-08 | ND | 9.6582e-10 | 2.5152e-07 | -- | Not Applicable | -- | -155.6%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 0.2 | ND | 0.00094016 | ND | 0.0010587 | ND | 3.4349e-08 | ND | 1.7174e-09 | ND | 4.4725e-07 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.05 | 0.05 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,1,1,2-Tetrachloroethane | Sample 1 | ND | ND | 0.05 | 0.05 | 1,1,1,2-Tetrachloroethane | 1 | 167.85 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00034913 | Not Applicable | 0.00034913 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00034913 | Not Applicable
| Background | ND | ND | 0.05 | 0.05 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Ethylbenzene | Sample 1 | 0.5 | 0.5 | -- | -- | Ethylbenzene | 1 | 106.17 | 0.5 | ND | 0.00088333 | ND | 0.00099251 | ND | 3.22e-08 | ND | 1.61e-09 | ND | 4.1927e-07 | ND | 1.6118e-09 | 4.1975e-07 | -- | Not Applicable | -- | 0%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 0.5 | ND | 0.00088333 | ND | 0.00099475 | ND | 3.2273e-08 | ND | 1.6136e-09 | ND | 4.2022e-07 | ND |  |  | -- | Not Applicable
| Background | 0.3 | 0.3 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
m&p-Xylene | Sample 1 | 1 | 2 | -- | -- | m&p-Xylene | 1 | 106 | 1 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | 0 | -- | Not Applicable | -- | -66.67%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 2 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 2 | 2 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
o-Xylene | Sample 1 | 2 | 2 | -- | -- | o-Xylene | 1 | 106 | 2 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | 0 | -- | Not Applicable | -- | 0%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 2 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 3 | 3 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Styrene | Sample 1 | 1 | 1 | -- | -- | Styrene | 1 | 104 | 1 | ND | 0.0043264 | ND | 0.0048611 | ND | 1.5771e-07 | ND | 7.8855e-09 | ND | 2.0535e-06 | ND | 7.8944e-09 | 2.0558e-06 | -- | Not Applicable | -- | 0%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 1 | ND | 0.0043264 | ND | 0.0048721 | ND | 1.5807e-07 | ND | 7.9033e-09 | ND | 2.0581e-06 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 0.07 | 0.07 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Bromoform | Sample 1 | ND | ND | 0.05 | 0.05 | Bromoform | 1 | 252.73 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00052568 | Not Applicable | 0.00052568 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00052568 | Not Applicable
| Background | ND | ND | 0.05 | 0.05 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Cumene | Sample 1 | ND | ND | 0.04 | 0.04 | Cumene | 1 | 120.19 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0002 | Not Applicable | 0.0002 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0002 | Not Applicable
| Background | ND | ND | 0.04 | 0.04 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,1,2,2-Tetrachloroethane | Sample 1 | ND | ND | 0.05 | 0.05 | 1,1,2,2-Tetrachloroethane | 1 | 168 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00034944 | Not Applicable | 0.00034944 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00034944 | Not Applicable
| Background | ND | ND | 0.05 | 0.05 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,2,3-Trichloropropane | Sample 1 | ND | ND | 0.08 | 0.08 | 1,2,3-Trichloropropane | 1 | 147.43 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00049065 | Not Applicable | 0.00049065 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00049065 | Not Applicable
| Background | ND | ND | 0.08 | 0.08 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Bromobenzene | Sample 1 | ND | ND | 0.06 | 0.06 | Bromobenzene | 1 | 157.01 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0003919 | Not Applicable | 0.0003919 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0003919 | Not Applicable
| Background | ND | ND | 0.06 | 0.06 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
4-Ethyltoluene | Sample 1 | 0.2 | 0.2 | -- | -- | 4-Ethyltoluene | 1 | 120.19 | 0.2 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | 0 | -- | Not Applicable | -- | 0%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 0.2 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 0.3 | 0.3 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,3,5-Trimethylbenzene | Sample 1 | ND | 0.3 | 0.05 | -- | 1,3,5-Trimethylbenzene | 1 | 120 | 0.025 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | 0 | -- | Not Applicable | -- | -169.2%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 0.3 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 0.3 | 0.3 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Alpha Methyl Styrene | Sample 1 | ND | ND | 0.06 | 0.06 | Alpha Methyl Styrene | 1 | 118.18 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00029498 | Not Applicable | 0.00029498 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00029498 | Not Applicable
| Background | ND | ND | 0.06 | 0.06 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,2,4-Trimethylbenzene | Sample 1 | 0.5 | 0.8 | -- | -- | 1,2,4-Trimethylbenzene | 1 | 120 | 0.5 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | 0 | -- | Not Applicable | -- | -46.15%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 0.8 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 0.9 | 0.9 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,3-Dichlorobenzene | Sample 1 | ND | ND | 0.07 | 0.07 | 1,3-Dichlorobenzene | 1 | 147 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00042806 | Not Applicable | 0.00042806 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00042806 | Not Applicable
| Background | ND | ND | 0.07 | 0.07 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,4-Dichlorobenzene | Sample 1 | ND | ND | 0.07 | 0.07 | 1,4-Dichlorobenzene | 1 | 147 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00042806 | Not Applicable | 0.00042806 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00042806 | Not Applicable
| Background | ND | ND | 0.07 | 0.07 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzylchloride | Sample 1 | ND | ND | 0.05 | 0.05 | Benzylchloride | 1 | 127 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00026416 | Not Applicable | 0.00026416 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00026416 | Not Applicable
| Background | ND | ND | 0.05 | 0.05 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,2-Dichlorobenzene | Sample 1 | ND | ND | 0.08 | 0.08 | 1,2-Dichlorobenzene | 1 | 147 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00048922 | Not Applicable | 0.00048922 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00048922 | Not Applicable
| Background | ND | ND | 0.08 | 0.08 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Hexachlorethane | Sample 1 | ND | ND | 0.06 | 0.06 | Hexachlorethane | 1 | 236.74 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0005909 | Not Applicable | 0.0005909 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0005909 | Not Applicable
| Background | ND | ND | 0.06 | 0.06 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,2,4-Trichlorobenzene | Sample 1 | ND | ND | 0.17 | 0.17 | 1,2,4-Trichlorobenzene | 1 | 181 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00128 | Not Applicable | 0.00128 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00128 | Not Applicable
| Background | ND | ND | 0.17 | 0.17 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Hexachlorobutadiene | Sample 1 | ND | ND | 0.07 | 0.07 | Hexachlorobutadiene | 1 | 261 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00076003 | Not Applicable | 0.00076003 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00076003 | Not Applicable
| Background | ND | ND | 0.07 | 0.07 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
VOC Tentatively Identified Compounds | [5 empty] | VOC Tentatively Identified Compounds
Pollutant | Sample ID | Sample Speciation |  | Detection Limits |  | Pollutant | Run Number | Molecular Weight (g/gmol) | Sample Speciation with Detection Limits (SSDL) (ppbv)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3)
|  | Run 1 | Run 2 | Run 1 | Run 2
|  | (ppbv) | (ppbv) | (ppbv) | (ppbv) |  |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
Ethanol | Sample 1 | ND | 1 | ND | -- | Ethanol | 1 | 46.0688 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0 | 0 | -- | Not Applicable | --
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 1 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 14 | 14 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
a If the corresponding sample speciation value does not = "ND" or all active sample speciation values = "ND," Sample Speciation with Detection Limits (SSDL) = Sample Speciation. Otherwise, (i.e., if the corresponding sample speciation value = "ND" and
at least one of the active sample speciation values does not = "ND," Sample Speciation with Detection Limits (SSDL) = 0.5 * Detection Limit.
b Background Corrected Concentration (BCC) = Test Concentration - Background Concentration
c Dilution-Corrected Concentration (DCC) = BCC / Dilution Correction Factor
d Quantity of Pollutant Emitted (QPE) =BDCC * Initial Plume Volume
e Emission Factor = QPE / No. of Items
f Emission Factor = QPE / Net Explosive Weight
g If Average Emission Factor = "ND", Minimum Detection Limit Based Concentration = Test Minimum DL in mg/m3

## Hydrocarbons
Hydrocarbon Data and Emission Factor Calculations
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
Test Run Number | 1 | 2 | Background
Number of Items Tested | 20 | 20 | 0
Net Explosive Weight for All Items (lb) | 0.0768 | 0.0768 | NA
Initial Plume Volume (m3) | 14.716 | 14.716 | 14.716
Dilution Correction Factor | 0.89 | 0.888 | NA
Is Sample 1 Active? | Yes | Yes | Yes
Is Sample 2 Active? | No | No | No
Pollutant | Sample ID | Sample Speciation |  | Detection Limits |  | Pollutant | Run Number | Molecular Weight (g/gmol) | Sample Speciation with Detection Limits (SSDL) (ppbv)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3) | Test Concentration Relative Percent Difference
|  | Run 1 | Run 2 | Run 1 | Run 2
|  | (ppbv) | (ppbv) | (ppbv) | (ppbv) |  |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
Methane | Sample 1 | 7100 | 6900 | -- | -- | Methane | 1 | 16.04 | 7100 | ND | 3.4698 | ND | 3.8986 | ND | 0.00012648 | ND | 6.3242e-06 | ND | 0.0016469 | ND | 6.2094e-06 | 0.001617 | -- | Not Applicable | -- | 2.857%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 6900 | ND | 3.3363 | ND | 3.7571 | ND | 0.00012189 | ND | 6.0946e-06 | ND | 0.0015871 | ND |  |  | -- | Not Applicable
| Background | 1900 | 1900 | -- | -- |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1-Butene | Sample 1 | ND | ND | 20 | 20 | 1-Butene | 1 | 56.11 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.046684 | Not Applicable | 0.046684 | ND
1-Butene | Sample 2 | ND | ND | ND | ND | 1-Butene | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.046684 | Not Applicable
1-Butene | Background | ND | ND | 20 | 20 | 1-Butene | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1-Butyne | Sample 1 | ND | ND | 20 | 20 | 1-Butyne | 1 | 54.09 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.045003 | Not Applicable | 0.045003 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.045003 | Not Applicable
| Background | ND | ND | 20 | 20 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2-Butyne | Sample 1 | ND | ND | 20 | 20 | 2-Butyne | 1 | 54.09 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.045003 | Not Applicable | 0.045003 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.045003 | Not Applicable
| Background | ND | ND | 20 | 20 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Acetylene | Sample 1 | 27 | 28 | -- | -- | Acetylene | 1 | 26.04 | 27 | ND | 0.029248 | ND | 0.032863 | ND | 1.0662e-06 | ND | 5.3309e-08 | ND | 1.3883e-05 | ND | 5.4359e-08 | 1.4156e-05 | -- | Not Applicable | -- | -3.636%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 28 | ND | 0.030331 | ND | 0.034157 | ND | 1.1082e-06 | ND | 5.5408e-08 | ND | 1.4429e-05 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 20 | 20 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,3-Butadiene | Sample 1 | ND | ND | 20 | 20 | 1,3-Butadiene | 1 | 54.09 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.045003 | Not Applicable | 0.045003 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.045003 | Not Applicable
| Background | ND | ND | 20 | 20 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Butane | Sample 1 | ND | ND | 22 | 22 | Butane | 1 | 58.12 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.053191 | Not Applicable | 0.053191 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.053191 | Not Applicable
| Background | ND | ND | 22 | 22 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
cis-butene | Sample 1 | ND | ND | 20 | 20 | cis-butene | 1 | 56.12 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.046692 | Not Applicable | 0.046692 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.046692 | Not Applicable
| Background | ND | ND | 20 | 20 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Ethane | Sample 1 | 87 | 84 | -- | -- | Ethane | 1 | 30.07 | 87 | ND | 0.10883 | ND | 0.12228 | ND | 3.9672e-06 | ND | 1.9836e-07 | ND | 5.1656e-05 | ND | 1.9515e-07 | 5.0821e-05 | -- | Not Applicable | -- | 3.509%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 84 | ND | 0.10508 | ND | 0.11833 | ND | 3.839e-06 | ND | 1.9195e-07 | ND | 4.9987e-05 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 20 | 20 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Ethylene | Sample 1 | 182 | 173 | -- | -- | Ethylene | 1 | 28.05 | 182 | ND | 0.21237 | ND | 0.23862 | ND | 7.7416e-06 | ND | 3.8708e-07 | ND | 0.0001008 | ND | 3.7792e-07 | 9.8418e-05 | -- | Not Applicable | -- | 5.07%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 173 | ND | 0.20187 | ND | 0.22733 | ND | 7.3754e-06 | ND | 3.6877e-07 | ND | 9.6033e-05 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 20 | 20 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
n-Hexane | Sample 1 | ND | ND | 20 | 20 | n-Hexane | 1 | 86.18 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.071702 | Not Applicable | 0.071702 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.071702 | Not Applicable
| Background | ND | ND | 20 | 20 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Isobutane | Sample 1 | ND | ND | 20 | 20 | Isobutane | 1 | 58.12 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.048356 | Not Applicable | 0.048356 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.048356 | Not Applicable
| Background | ND | ND | 20 | 20 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Isobutylene | Sample 1 | ND | ND | 20 | 20 | Isobutylene | 1 | 56.1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.046675 | Not Applicable | 0.046675 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.046675 | Not Applicable
| Background | ND | ND | 20 | 20 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
n-Pentane | Sample 1 | ND | ND | 20 | 20 | n-Pentane | 1 | 72.15 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.060029 | Not Applicable | 0.060029 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.060029 | Not Applicable
| Background | ND | ND | 20 | 20 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Propane | Sample 1 | ND | ND | 20 | 20 | Propane | 1 | 44.1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.036691 | Not Applicable | 0.036691 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.036691 | Not Applicable
| Background | ND | ND | 20 | 20 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Propylene | Sample 1 | 23 | ND | -- | 20 | Propylene | 1 | 42.08 | 23 | ND | 0.040262 | ND | 0.045238 | ND | 1.4677e-06 | ND | 7.3384e-08 | ND | 1.911e-05 | ND | 5.2681e-08 | 1.3719e-05 | -- | Not Applicable | -- | 78.79%
| Sample 2 | ND | ND | ND | ND |  | 2 |  | 10 | ND | 0.017505 | ND | 0.019713 | ND | 6.3956e-07 | ND | 3.1978e-08 | ND | 8.3276e-06 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 20 | 20 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Propyne | Sample 1 | ND | ND | 20 | 20 | Propyne | 1 | 39.1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.032531 | Not Applicable | 0.032531 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.032531 | Not Applicable
| Background | ND | ND | 20 | 20 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
trans-Butene | Sample 1 | ND | ND | 20 | 20 | trans-Butene | 1 | 56.12 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.046692 | Not Applicable | 0.046692 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 |  | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.046692 | Not Applicable
| Background | ND | ND | 20 | 20 |  | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
a If the corresponding sample speciation value does not = "ND" or all active sample speciation values = "ND," Sample Speciation with Detection Limits (SSDL) = Sample Speciation. Otherwise, (i.e., if the corresponding sample speciation value = "ND" and
at least one of the active sample speciation values does not = "ND," Sample Speciation with Detection Limits (SSDL) = 0.5 * Detection Limit.
b Background Corrected Concentration (BCC) = Test Concentration - Background Concentration
c Dilution-Corrected Concentration (DCC) = BCC / Dilution Correction Factor
d Quantity of Pollutant Emitted (QPE) =BDCC * Initial Plume Volume
e Emission Factor = QPE / No. of Items
f Emission Factor = QPE / Net Explosive Weight
g If Average Emission Factor = "ND", Minimum Detection Limit Based Concentration = Test Minimum DL in mg/m3

## SVOC & SVOC TICs
SVOC Data and Emission Factor Calculations
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
Test Run Number | 1 | 2 | Background
Number of Items Tested | 20 | 20 | 0
Net Explosive Weight for All Items (lb) | 0.0768 | 0.0768 | NA
Initial Plume Volume (m3) | 14.716 | 14.716 | 14.716
Dilution Correction Factor | 0.89 | 0.888 | NA
Is Sample 1 Active? | Yes | Yes | Yes
Is Sample 2 Active? | No | No | No
SVOC Sample Volume (m3) Corrected to EPA Stationary Source Program Standard Conditions
Sample 1 | 0.61475 | 0.59989
Sample 2 | NA | NA
Background | 0.58147 | 0.58147
Pollutant | Sample ID | Sample Speciation (ug) |  | Detection Limits (ug) |  | Pollutant | Run Number | Sample Speciation with Detection Limits (SSDL) (ug)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC) (mg/m3)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3) | Test Concentration Relative Percent Difference
|  | Run 1 | Run 2 | Run 1 | Run 2 |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
N-nitrosodimethylamine | Sample 1 | ND | ND | 10 | 10 | N-nitrosodimethylamine | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Bis(2-chloroethyl)ether | Sample 1 | ND | ND | 10 | 10 | Bis(2-chloroethyl)ether | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Phenol | Sample 1 | ND | ND | 10 | 10 | Phenol | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2-chlorophenol | Sample 1 | ND | ND | 10 | 10 | 2-chlorophenol | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,3-dichlorobenzene | Sample 1 | ND | ND | 10 | 10 | 1,3-dichlorobenzene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,4-dichlorobenzene | Sample 1 | ND | ND | 10 | 10 | 1,4-dichlorobenzene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,2-dichlorobenzene | Sample 1 | ND | ND | 10 | 10 | 1,2-dichlorobenzene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzyl alcohol | Sample 1 | ND | ND | 10 | 10 | Benzyl alcohol | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Bis(2-chloroisopropyl)ether | Sample 1 | ND | ND | 10 | 10 | Bis(2-chloroisopropyl)ether | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2-methylphenol | Sample 1 | ND | ND | 10 | 10 | 2-methylphenol | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Hexachloroethane | Sample 1 | ND | ND | 10 | 10 | Hexachloroethane | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
N-nitroso-di-n-propylamine | Sample 1 | ND | ND | 10 | 10 | N-nitroso-di-n-propylamine | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
4-methylphenol | Sample 1 | ND | ND | 10 | 10 | 4-methylphenol | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Nitrobenzene | Sample 1 | ND | ND | 10 | 10 | Nitrobenzene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Isophorone | Sample 1 | ND | ND | 10 | 10 | Isophorone | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2-nitrophenol | Sample 1 | ND | ND | 10 | 10 | 2-nitrophenol | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,4-dimethylphenol | Sample 1 | ND | ND | 10 | 10 | 2,4-dimethylphenol | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Bis(2-chloroethoxy)methane | Sample 1 | ND | ND | 10 | 10 | Bis(2-chloroethoxy)methane | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,4-dichlorophenol | Sample 1 | ND | ND | 10 | 10 | 2,4-dichlorophenol | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,2,4-trichlorobenzene | Sample 1 | ND | ND | 10 | 10 | 1,2,4-trichlorobenzene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Naphthalene | Sample 1 | ND | ND | 10 | 10 | Naphthalene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
4-chloroaniline | Sample 1 | ND | ND | 10 | 10 | 4-chloroaniline | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Hexachlorobutadiene | Sample 1 | ND | ND | 10 | 10 | Hexachlorobutadiene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
4-chloro-3-methylphenol | Sample 1 | ND | ND | 10 | 10 | 4-chloro-3-methylphenol | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2-Methylnaphthalene | Sample 1 | ND | ND | 10 | 10 | 2-Methylnaphthalene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Hexachlorocyclopentadiene | Sample 1 | ND | ND | 10 | 10 | Hexachlorocyclopentadiene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,4,6-trichlorophenol | Sample 1 | ND | ND | 10 | 10 | 2,4,6-trichlorophenol | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,4,5-trichlorophenol | Sample 1 | ND | ND | 10 | 10 | 2,4,5-trichlorophenol | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2-chloronaphthalene | Sample 1 | ND | ND | 10 | 10 | 2-chloronaphthalene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2-nitroaniline | Sample 1 | ND | ND | 10 | 10 | 2-nitroaniline | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Acenaphthylene | Sample 1 | ND | ND | 10 | 10 | Acenaphthylene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Dimethylphthalate | Sample 1 | ND | ND | 10 | 10 | Dimethylphthalate | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,6-dinitrotoluene | Sample 1 | ND | ND | 10 | 10 | 2,6-dinitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Acenaphthene | Sample 1 | ND | ND | 10 | 10 | Acenaphthene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
3-nitroaniline | Sample 1 | ND | ND | 10 | 10 | 3-nitroaniline | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,4-dinitrophenol | Sample 1 | ND | ND | 10 | 10 | 2,4-dinitrophenol | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Dibenzofuran | Sample 1 | ND | ND | 10 | 10 | Dibenzofuran | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,4-dinitrotoluene | Sample 1 | ND | ND | 10 | 10 | 2,4-dinitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
4-nitrophenol | Sample 1 | ND | ND | 10 | 10 | 4-nitrophenol | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Fluorene | Sample 1 | ND | ND | 10 | 10 | Fluorene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
4-chlorophenyl-phenylether | Sample 1 | ND | ND | 10 | 10 | 4-chlorophenyl-phenylether | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Diethylphthalate | Sample 1 | ND | ND | 10 | 10 | Diethylphthalate | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
4-nitroaniline | Sample 1 | ND | ND | 10 | 10 | 4-nitroaniline | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
4,6-dinitro-2-methylphenol | Sample 1 | ND | ND | 10 | 10 | 4,6-dinitro-2-methylphenol | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
N-nitrosodiphenylamine(1) | Sample 1 | ND | ND | 10 | 10 | N-nitrosodiphenylamine(1) | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
4-bromophenyl-phenylether | Sample 1 | ND | ND | 10 | 10 | 4-bromophenyl-phenylether | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Hexachlorobenzene | Sample 1 | ND | ND | 10 | 10 | Hexachlorobenzene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Pentachlorophenol | Sample 1 | ND | ND | 10 | 10 | Pentachlorophenol | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Phenanthrene | Sample 1 | ND | ND | 10 | 10 | Phenanthrene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Anthracene | Sample 1 | ND | ND | 10 | 10 | Anthracene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Di-n-butylphthalate | Sample 1 | ND | ND | 10 | 10 | Di-n-butylphthalate | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Fluoranthene | Sample 1 | ND | ND | 10 | 10 | Fluoranthene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Pyrene | Sample 1 | ND | ND | 10 | 10 | Pyrene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Butylbenzylphthalate | Sample 1 | ND | ND | 10 | 10 | Butylbenzylphthalate | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzo(a)anthracene | Sample 1 | ND | ND | 10 | 10 | Benzo(a)anthracene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Chrysene | Sample 1 | ND | ND | 10 | 10 | Chrysene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
bis(2-Ethylhexyl)phthalate | Sample 1 | 15 | ND | -- | 10 | bis(2-Ethylhexyl)phthalate | 1 | 15 | ND | 0.0244 | ND | 0.027416 | ND | 8.8945e-07 | ND | 4.4473e-08 | ND | 1.1581e-05 | ND | 2.9849e-08 | 7.7732e-06 | -- | Not Applicable | -- | 98.15%
| Sample 2 | ND | ND | ND | ND |  | 2 | 5 | ND | 0.0083349 | ND | 0.0093862 | ND | 3.0452e-07 | ND | 1.5226e-08 | ND | 3.9651e-06 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Di-n-octylphthalate | Sample 1 | ND | ND | 10 | 10 | Di-n-octylphthalate | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzo(b)fluoranthene | Sample 1 | ND | ND | 10 | 10 | Benzo(b)fluoranthene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzo(k)fluoranthene | Sample 1 | ND | ND | 10 | 10 | Benzo(k)fluoranthene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzo(a)pyrene | Sample 1 | ND | ND | 10 | 10 | Benzo(a)pyrene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Indeno(1,2,3-cd)pyrene | Sample 1 | ND | ND | 10 | 10 | Indeno(1,2,3-cd)pyrene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Dibenz(a,h)anthracene | Sample 1 | ND | ND | 10 | 10 | Dibenz(a,h)anthracene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzo(g,h,i)perylene | Sample 1 | ND | ND | 10 | 10 | Benzo(g,h,i)perylene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.016267 | Not Applicable | 0.016267 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.01667 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
SVOC Tentatively Identified Compounds | [5 empty] | SVOC Tentatively Identified Compounds
Pollutant | Sample ID | Sample Speciation (ug) |  | Detection Limits (ug) |  | Pollutant | Run Number | Sample Speciation with Detection Limits (SSDL) (ug)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC) (mg/m3)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3)
|  | Run 1 | Run 2 | Run 1 | Run 2 |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
1-Butanol, 3-methyl-, acetate | Sample 1 | 16 | 24 | -- | -- | 1-Butanol, 3-methyl-, acetate | 1 | 16 | ND | 0.0019496 | ND | 0.0021906 | ND | 7.107e-08 | ND | 3.5535e-09 | ND | 9.2539e-07 | ND | 1.6327e-08 | 4.2519e-06 | -- | Not Applicable | --
| Sample 2 | ND | ND | ND | ND |  | 2 | 24 | ND | 0.015931 | ND | 0.01794 | ND | 5.8202e-07 | ND | 2.9101e-08 | ND | 7.5784e-06 | ND |  |  | -- | Not Applicable
| Background | 14 | 14 | -- | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Aectophenone | Sample 1 | 11 | ND | -- | ND | Aectophenone | 1 | 11 | ND | 0.017893 | ND | 0.020105 | ND | 6.5227e-07 | ND | 3.2613e-08 | ND | 8.493e-06 | ND | 3.2613e-08 | 8.493e-06 | -- | Not Applicable | --
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | -- | Not Applicable
| Background | ND | ND | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Isophthalaldehyde | Sample 1 | 12 | ND | -- | ND | Isophthalaldehyde | 1 | 12 | ND | 0.01952 | ND | 0.021933 | ND | 7.1156e-07 | ND | 3.5578e-08 | ND | 9.2651e-06 | ND | 3.5578e-08 | 9.2651e-06 | -- | Not Applicable | --
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | -- | Not Applicable
| Background | ND | ND | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzaldehyde, 4-ethyl- | Sample 1 | 7.6 | ND | -- | ND | Benzaldehyde, 4-ethyl- | 1 | 7.6 | ND | 0.012363 | ND | 0.013891 | ND | 4.5066e-07 | ND | 2.2533e-08 | ND | 5.8679e-06 | ND | 2.2533e-08 | 5.8679e-06 | -- | Not Applicable | --
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | -- | Not Applicable
| Background | ND | ND | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Eicosane | Sample 1 | 4.6 | ND | -- | ND | Eicosane | 1 | 4.6 | ND | 0.0074827 | ND | 0.0084075 | ND | 2.7277e-07 | ND | 1.3638e-08 | ND | 3.5516e-06 | ND | 1.3638e-08 | 3.5516e-06 | -- | Not Applicable | --
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | -- | Not Applicable
| Background | ND | ND | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
a If the corresponding sample speciation value does not = "ND" or all active sample speciation values = "ND," Sample Speciation with Detection Limits (SSDL) = Sample Speciation. Otherwise, (i.e., if the corresponding sample speciation value = "ND" and
at least one of the active sample speciation values does not = "ND," Sample Speciation with Detection Limits (SSDL) = 0.5 * Detection Limit.
b Background Corrected Concentration (BCC) = Test Mass/Test Sample Volume - Background Mass/ Background Sample Volume
c Dilution-Corrected Concentration (DCC) = BCC / Dilution Correction Factor
d Quantity of Pollutant Emitted (QPE) =BDCC * Initial Plume Volume
e Emission Factor = QPE / No. of Items
f Emission Factor = QPE / Net Explosive Weight
g If Average Emission Factor = "ND", Minimum Detection Limit Based Concentration = Test DL/Test Sample volume in mg/m3

## PAH
TO-13 PAH Data and Emission Factor Calculations
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
Test Run Number | 1 | 2 | Background
Number of Items Tested | 20 | 20 | 0
Net Explosive Weight for All Items (lb) | 0.0768 | 0.0768 | NA
Initial Plume Volume (m3) | 14.716 | 14.716 | 14.716
Dilution Correction Factor | 0.89 | 0.888 | NA
Is Sample 1 Active? | Yes | Yes | Yes
Is Sample 2 Active? | No | No | No
TO-13_PAH Sample Volume (m3) Corrected to EPA Stationary Source Program Standard Conditions
Sample 1 | 0.61475 | 0.59989
Sample 2 | NA | NA
Background | 0.58147 | 0.58147
Pollutant | Sample ID | Sample Speciation |  | Detection Limits |  | Pollutant | Run Number | Sample Speciation with Detection Limits (SSDL) (ng)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC) (mg/m3)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3) | Test Concentration Relative Percent Difference
|  | Run 1 | Run 2 | Run 1 | Run 2
|  | (ng) | (ng) | (ng) | (ng) |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
Naphthalene | Sample 1 | 4600 | 4200 | -- | -- | Naphthalene | 1 | 4600 | ND | 0.0072075 | ND | 0.0080983 | ND | 2.6273e-07 | ND | 1.3137e-08 | ND | 3.421e-06 | ND | 1.2712e-08 | 3.3104e-06 | -- | Not Applicable | -- | 6.646%
| Sample 2 | ND | ND | ND | ND |  | 2 | 4200 | ND | 0.0067262 | ND | 0.0075745 | ND | 2.4574e-07 | ND | 1.2287e-08 | ND | 3.1998e-06 | ND |  |  | -- | Not Applicable
| Background | 160 | 160 | -- | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Acenaphthylene | Sample 1 | 420 | 380 | -- | -- | Acenaphthylene | 1 | 420 | ND | 0.0006832 | ND | 0.00076764 | ND | 2.4905e-08 | ND | 1.2452e-09 | ND | 3.2428e-07 | ND | 1.2012e-09 | 3.1281e-07 | -- | Not Applicable | -- | 7.556%
| Sample 2 | ND | ND | ND | ND |  | 2 | 380 | ND | 0.00063345 | ND | 0.00071335 | ND | 2.3143e-08 | ND | 1.1572e-09 | ND | 3.0134e-07 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Acenaphthene | Sample 1 | 90 | 82 | -- | -- | Acenaphthene | 1 | 90 | ND | 9.9966e-05 | ND | 0.00011232 | ND | 3.644e-09 | ND | 1.822e-10 | ND | 4.7449e-08 | ND | 1.7354e-10 | 4.5193e-08 | -- | Not Applicable | -- | 6.858%
| Sample 2 | ND | ND | ND | ND |  | 2 | 82 | ND | 9.0258e-05 | ND | 0.00010164 | ND | 3.2976e-09 | ND | 1.6488e-10 | ND | 4.2937e-08 | ND |  |  | -- | Not Applicable
| Background | 27 | 27 | -- | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Fluorene | Sample 1 | 180 | 170 | -- | -- | Fluorene | 1 | 180 | ND | 0.00024637 | ND | 0.00027682 | ND | 8.9808e-09 | ND | 4.4904e-10 | ND | 1.1694e-07 | ND | 4.4095e-10 | 1.1483e-07 | -- | Not Applicable | -- | 3.267%
| Sample 2 | ND | ND | ND | ND |  | 2 | 170 | ND | 0.00023695 | ND | 0.00026684 | ND | 8.6571e-09 | ND | 4.3285e-10 | ND | 1.1272e-07 | ND |  |  | -- | Not Applicable
| Background | 27 | 27 | -- | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Phenanthrene | Sample 1 | 250 | 200 | -- | -- | Phenanthrene | 1 | 250 | ND | 0.00029316 | ND | 0.00032939 | ND | 1.0687e-08 | ND | 5.3433e-10 | ND | 1.3915e-07 | ND | 4.6801e-10 | 1.2188e-07 | -- | Not Applicable | -- | 19.8%
| Sample 2 | ND | ND | ND | ND |  | 2 | 200 | ND | 0.00021989 | ND | 0.00024762 | ND | 8.0337e-09 | ND | 4.0169e-10 | ND | 1.0461e-07 | ND |  |  | -- | Not Applicable
| Background | 66 | 66 | -- | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Anthracene | Sample 1 | 40 | 21 | -- | -- | Anthracene | 1 | 40 | ND | 6.5067e-05 | ND | 7.3109e-05 | ND | 2.3719e-09 | ND | 1.1859e-10 | ND | 3.0884e-08 | ND | 9.1271e-11 | 2.3769e-08 | -- | Not Applicable | -- | 60.08%
| Sample 2 | ND | ND | ND | ND |  | 2 | 21 | ND | 3.5007e-05 | ND | 3.9422e-05 | ND | 1.279e-09 | ND | 6.3949e-11 | ND | 1.6653e-08 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Fluoranthene | Sample 1 | 130 | 26 | -- | -- | Fluoranthene | 1 | 130 | ND | 0.00018567 | ND | 0.00020862 | ND | 6.7682e-09 | ND | 3.3841e-10 | ND | 8.8128e-08 | ND | 1.8523e-10 | 4.8237e-08 | -- | Not Applicable | -- | 132%
| Sample 2 | ND | ND | ND | ND |  | 2 | 26 | ND | 1.7545e-05 | ND | 1.9758e-05 | ND | 6.41e-10 | ND | 3.205e-11 | ND | 8.3463e-09 | ND |  |  | -- | Not Applicable
| Background | 15 | 15 | -- | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Pyrene | Sample 1 | 270 | 35 | -- | -- | Pyrene | 1 | 270 | ND | 0.00041168 | ND | 0.00046256 | ND | 1.5007e-08 | ND | 7.5035e-10 | ND | 1.954e-07 | ND | 4.0333e-10 | 1.0504e-07 | -- | Not Applicable | -- | 153.1%
| Sample 2 | ND | ND | ND | ND |  | 2 | 35 | ND | 3.0828e-05 | ND | 3.4716e-05 | ND | 1.1263e-09 | ND | 5.6315e-11 | ND | 1.4665e-08 | ND |  |  | -- | Not Applicable
| Background | 16 | 16 | -- | -- |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzo(a)anthracene | Sample 1 | ND | ND | 50 | 50 | Benzo(a)anthracene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 8.1333e-05 | Not Applicable | 8.1333e-05 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 8.3349e-05 | Not Applicable
| Background | ND | ND | 50 | 50 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Chrysene | Sample 1 | ND | ND | 50 | 50 | Chrysene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 8.1333e-05 | Not Applicable | 8.1333e-05 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 8.3349e-05 | Not Applicable
| Background | ND | ND | 50 | 50 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzo(b)fluoranthene | Sample 1 | 56 | ND | -- | 50 | Benzo(b)fluoranthene | 1 | 56 | ND | 9.1093e-05 | ND | 0.00010235 | ND | 3.3206e-09 | ND | 1.6603e-10 | ND | 4.3237e-08 | ND | 1.2108e-10 | 3.1531e-08 | -- | Not Applicable | -- | 74.44%
| Sample 2 | ND | ND | ND | ND |  | 2 | 25 | ND | 4.1675e-05 | ND | 4.6931e-05 | ND | 1.5226e-09 | ND | 7.6129e-11 | ND | 1.9825e-08 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 50 | 50 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzo(k)fluoranthene | Sample 1 | ND | ND | 50 | 50 | Benzo(k)fluoranthene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 8.1333e-05 | Not Applicable | 8.1333e-05 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 8.3349e-05 | Not Applicable
| Background | ND | ND | 50 | 50 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzo(e)pyrene | Sample 1 | 57 | ND | -- | 50 | Benzo(e)pyrene | 1 | 57 | ND | 9.272e-05 | ND | 0.00010418 | ND | 3.3799e-09 | ND | 1.69e-10 | ND | 4.4009e-08 | ND | 1.2256e-10 | 3.1917e-08 | -- | Not Applicable | -- | 75.96%
| Sample 2 | ND | ND | ND | ND |  | 2 | 25 | ND | 4.1675e-05 | ND | 4.6931e-05 | ND | 1.5226e-09 | ND | 7.6129e-11 | ND | 1.9825e-08 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 50 | 50 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzo(a)pyrene | Sample 1 | 55 | ND | -- | 50 | Benzo(a)pyrene | 1 | 55 | ND | 8.9467e-05 | ND | 0.00010052 | ND | 3.2613e-09 | ND | 1.6307e-10 | ND | 4.2465e-08 | ND | 1.196e-10 | 3.1145e-08 | -- | Not Applicable | -- | 72.89%
| Sample 2 | ND | ND | ND | ND |  | 2 | 25 | ND | 4.1675e-05 | ND | 4.6931e-05 | ND | 1.5226e-09 | ND | 7.6129e-11 | ND | 1.9825e-08 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 50 | 50 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Indeno(1,2,3-cd)pyrene | Sample 1 | 50 | ND | -- | 10 | Indeno(1,2,3-cd)pyrene | 1 | 50 | ND | 8.1333e-05 | ND | 9.1386e-05 | ND | 2.9648e-09 | ND | 1.4824e-10 | ND | 3.8605e-08 | ND | 8.1734e-11 | 2.1285e-08 | -- | Not Applicable | -- | 162.8%
| Sample 2 | ND | ND | ND | ND |  | 2 | 5 | ND | 8.3349e-06 | ND | 9.3862e-06 | ND | 3.0452e-10 | ND | 1.5226e-11 | ND | 3.9651e-09 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Dibenz(a,h)anthracene | Sample 1 | ND | ND | 10 | 10 | Dibenz(a,h)anthracene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 1.6267e-05 | Not Applicable | 1.6267e-05 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 1.667e-05 | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Benzo(g,h,i)perylene | Sample 1 | 130 | ND | -- | 10 | Benzo(g,h,i)perylene | 1 | 130 | ND | 0.00021147 | ND | 0.0002376 | ND | 7.7086e-09 | ND | 3.8543e-10 | ND | 1.0037e-07 | ND | 2.0033e-10 | 5.2169e-08 | -- | Not Applicable | -- | 184.8%
| Sample 2 | ND | ND | ND | ND |  | 2 | 5 | ND | 8.3349e-06 | ND | 9.3862e-06 | ND | 3.0452e-10 | ND | 1.5226e-11 | ND | 3.9651e-09 | ND |  |  | -- | Not Applicable
| Background | ND | ND | 10 | 10 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
a If the corresponding sample speciation value does not = "ND" or all active sample speciation values = "ND," Sample Speciation with Detection Limits (SSDL) = Sample Speciation. Otherwise, (i.e., if the corresponding sample speciation value = "ND" and
at least one of the active sample speciation values does not = "ND," Sample Speciation with Detection Limits (SSDL) = 0.5 * Detection Limit.
b Background Corrected Concentration (BCC) = Test Mass/Test Sample Volume - Background Mass/ Background Sample Volume
c Dilution-Corrected Concentration (DCC) = BCC / Dilution Correction Factor
d Quantity of Pollutant Emitted (QPE) =BDCC * Initial Plume Volume
e Emission Factor = QPE / No. of Items
f Emission Factor = QPE / Net Explosive Weight
g If Average Emission Factor = "ND", Minimum Detection Limit Based Concentration = Test DL/Test Sample volume in mg/m3
Note on A39: MACTEC Employee: Present in the method blank at a level greater than 50% of test values.

## Dioxin-Furan
Dioxin / Furan Data and Emission Factor Calculations
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
Test Run Number | 1 | 2 | Background
Number of Items Tested | 20 | 20 | 0
Net Explosive Weight for All Items (lb) | 0.0768 | 0.0768 | NA
Initial Plume Volume (m3) | 14.716 | 14.716 | 14.716
Dilution Correction Factor | 0.89 | 0.888 | NA
Is Sample 1 Active? | Yes | Yes | Yes
Is Sample 2 Active? | No | No | No
Analytical Data Corrected for Volume at EPA Stationary Source Program Standard Conditions
Pollutant | Sample ID | Sample Speciation (pg/m3) |  | Detection Limits (pg/m3) |  | Pollutant | Run Number | Sample Speciation with Detection Limits (SSDL) (pg/m3)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC) (mg/m3)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3) | Test Concentration Relative Percent Difference
|  | Run 1 | Run 2 | Run 1 | Run 2 |  |  | Train A | Train B | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Train A | Train B | Train A | Train B | (lb/item) | (lb/lb NEW) | Train A | Train B
2378-TCDD | Sample 1 | 26.056 | 23.271 | ND | ND | 2378-TCDD | 1 | 26.056 | ND | 1.7374e-08 | ND | 1.9522e-08 | ND | 6.3334e-13 | ND | 3.1667e-14 | ND | 8.2467e-12 | ND | 2.9159e-14 | 7.5935e-12 | -- | Not Applicable | -- | 11.29%
| Sample 2 | ND | ND | ND | ND |  | 2 | 23.271 | ND | 1.4589e-08 | ND | 1.6429e-08 | ND | 5.3302e-13 | ND | 2.6651e-14 | ND | 6.9404e-12 | ND |  |  | -- | Not Applicable
| Background | 8.6817 | 8.6817 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
12378-PECDD | Sample 1 | 7.2239 | 2.2446 | ND | ND | 12378-PECDD | 1 | 7.2239 | ND | 4.4223e-09 | ND | 4.9689e-09 | ND | 1.6121e-13 | ND | 8.0603e-15 | ND | 2.099e-12 | ND | 4.0302e-15 | 1.0495e-12 | -- | Not Applicable | -- | 105.2%
| Sample 2 | ND | ND | ND | ND |  | 2 | 2.2446 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 2.8015 | 2.8015 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
123478-HXCDD | Sample 1 | 4.7393 | 1.3155 | ND | ND | 123478-HXCDD | 1 | 4.7393 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | 0 | -- | Not Applicable | -- | 113.1%
| Sample 2 | ND | ND | ND | ND |  | 2 | 1.3155 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 9.7182 | 9.7182 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
123678-HXCDD | Sample 1 | 26.317 | 8.1015 | ND | ND | 123678-HXCDD | 1 | 26.317 | ND | 1.1139e-08 | ND | 1.2516e-08 | ND | 4.0606e-13 | ND | 2.0303e-14 | ND | 5.2873e-12 | ND | 1.0152e-14 | 2.6436e-12 | -- | Not Applicable | -- | 105.8%
| Sample 2 | ND | ND | ND | ND |  | 2 | 8.1015 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 15.1776 | 15.1776 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
123789-HXCDD | Sample 1 | 11.8797 | 3.9777 | ND | ND | 123789-HXCDD | 1 | 11.8797 | ND | 6.8036e-11 | ND | 7.6445e-11 | ND | 2.4801e-15 | ND | 1.2401e-16 | ND | 3.2293e-14 | ND | 6.2003e-17 | 1.6147e-14 | -- | Not Applicable | -- | 99.66%
| Sample 2 | ND | ND | ND | ND |  | 2 | 3.9777 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 11.8117 | 11.8117 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,2,3,4,6,7,8-HpCDD | Sample 1 | 137.2637 | 40.9669 | ND | ND | 1,2,3,4,6,7,8-HpCDD | 1 | 137.2637 | ND | 4.4279e-08 | ND | 4.9752e-08 | ND | 1.6141e-12 | ND | 8.0705e-14 | ND | 2.1017e-11 | ND | 4.0353e-14 | 1.0508e-11 | -- | Not Applicable | -- | 108.1%
| Sample 2 | ND | ND | ND | ND |  | 2 | 40.9669 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 92.9847 | 92.9847 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
OCDD | Sample 1 | 359.8043 | 144.8145 | ND | ND | OCDD | 1 | 359.8043 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | 0 | -- | Not Applicable | -- | 85.21%
| Sample 2 | ND | ND | ND | ND |  | 2 | 144.8145 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 810.5298 | 810.5298 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,3,7,8-TCDF | Sample 1 | 32.5407 | 10.3837 | ND | ND | 2,3,7,8-TCDF | 1 | 32.5407 | ND | 2.289e-08 | ND | 2.5719e-08 | ND | 8.3442e-13 | ND | 4.1721e-14 | ND | 1.0865e-11 | ND | 2.153e-14 | 5.6068e-12 | -- | Not Applicable | -- | 103.2%
| Sample 2 | ND | ND | ND | ND |  | 2 | 10.3837 | ND | 7.3326e-10 | ND | 8.2574e-10 | ND | 2.679e-14 | ND | 1.3395e-15 | ND | 3.4882e-13 | ND |  |  | -- | Not Applicable
| Background | 9.6505 | 9.6505 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
12378-PECDF | Sample 1 | 14.9279 | 4.5936 | ND | ND | 12378-PECDF | 1 | 14.9279 | ND | 1.0187e-08 | ND | 1.1446e-08 | ND | 3.7134e-13 | ND | 1.8567e-14 | ND | 4.8352e-12 | ND | 9.2835e-15 | 2.4176e-12 | -- | Not Applicable | -- | 105.9%
| Sample 2 | ND | ND | ND | ND |  | 2 | 4.5936 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 4.7411 | 4.7411 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
23478-PECDF | Sample 1 | 15.596 | 4.0612 | ND | ND | 23478-PECDF | 1 | 15.596 | ND | 1.1183e-08 | ND | 1.2566e-08 | ND | 4.0767e-13 | ND | 2.0383e-14 | ND | 5.3081e-12 | ND | 1.0192e-14 | 2.6541e-12 | -- | Not Applicable | -- | 117.4%
| Sample 2 | ND | ND | ND | ND |  | 2 | 4.0612 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 4.4127 | 4.4127 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,2,3,4,7,8-HxCDF | Sample 1 | 23.1435 | 7.4751 | ND | ND | 1,2,3,4,7,8-HxCDF | 1 | 23.1435 | ND | 1.0993e-08 | ND | 1.2352e-08 | ND | 4.0073e-13 | ND | 2.0037e-14 | ND | 5.2179e-12 | ND | 1.0018e-14 | 2.6089e-12 | -- | Not Applicable | -- | 102.3%
| Sample 2 | ND | ND | ND | ND |  | 2 | 7.4751 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 12.1503 | 12.1503 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,2,3,6,7,8-HxCDF | Sample 1 | 9.6144 | 3.4557 | ND | ND | 1,2,3,6,7,8-HxCDF | 1 | 9.6144 | ND | 2.667e-09 | ND | 2.9966e-09 | ND | 9.7219e-14 | ND | 4.861e-15 | ND | 1.2659e-12 | ND | 2.4305e-15 | 6.3294e-13 | -- | Not Applicable | -- | 94.24%
| Sample 2 | ND | ND | ND | ND |  | 2 | 3.4557 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 6.9474 | 6.9474 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
123789-HXCDF | Sample 1 | ND | ND | 14.4477 | 3.6854 | 123789-HXCDF | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 1.4448e-08 | Not Applicable | 3.6854e-09 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 3.6854e-09 | Not Applicable
| Background | 7.7581 | 7.7581 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,3,4,6,7,8-HxCDF | Sample 1 | 5.2613 | 1.7853 | ND | ND | 2,3,4,6,7,8-HxCDF | 1 | 5.2613 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | 0 | -- | Not Applicable | -- | 98.66%
| Sample 2 | ND | ND | ND | ND |  | 2 | 1.7853 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 10.416 | 10.416 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,2,3,4,6,7,8-HpCDF | Sample 1 | 82.8655 | 25.1397 | ND | ND | 1,2,3,4,6,7,8-HpCDF | 1 | 82.8655 | ND | 3.2592e-08 | ND | 3.662e-08 | ND | 1.1881e-12 | ND | 5.9403e-14 | ND | 1.547e-11 | ND | 2.9702e-14 | 7.7348e-12 | -- | Not Applicable | -- | 106.9%
| Sample 2 | ND | ND | ND | ND |  | 2 | 25.1397 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 50.2739 | 50.2739 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1234789-HPCDF | Sample 1 | 2.4219 | 2.4221 | ND | ND | 1234789-HPCDF | 1 | 2.4219 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | 0 | -- | Not Applicable | -- | -0.009566%
| Sample 2 | ND | ND | ND | ND |  | 2 | 2.4221 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 40.5865 | 40.5865 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
OCDF | Sample 1 | 18.6547 | 14.4073 | ND | ND | OCDF | 1 | 18.6547 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | 0 | -- | Not Applicable | -- | 25.69%
| Sample 2 | ND | ND | ND | ND |  | 2 | 14.4073 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND | 0 | ND |  |  | -- | Not Applicable
| Background | 246.9467 | 246.9467 | ND | ND |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
[6 empty] | Total Dioxinh | 1 | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | 1.6691e-13 | 4.3466e-11 | -- | -- | -- | --
a If the corresponding sample speciation value does not = "ND" or all active sample speciation values = "ND," Sample Speciation with Detection Limits (SSDL) = Sample Speciation. Otherwise, (i.e., if the corresponding sample speciation value = "ND" and
at least one of the active sample speciation values does not = "ND," Sample Speciation with Detection Limits (SSDL) = 0.5 * Detection Limit.
b Background Corrected Concentration (BCC) = Test Concentration - Background Concentration
c Dilution-Corrected Concentration (DCC) = BCC / Dilution Correction Factor
d Quantity of Pollutant Emitted (QPE) =BDCC * Initial Plume Volume
e Emission Factor = QPE / No. of Items
Test Run Number | Run 1 | Run 2 | Bkgd |  |  | f Emission Factor = QPE / Net Explosive Weight
Volume at Actual Conditions (m3) | 0.65013 | 0.63298 | 0.61171 |  |  | g If Average Emission Factor = "ND", Minimum Detection Limit Based Concentration = Test Minimum DL in mg/m3
Volume at EPA Stationary Source Program Standard Conditions (m3) | 0.62279 | 0.6063 | 0.59608 |  |  | h Total Dioxin = Sum of all emission factors of dioxin compounds included in AP-42.
Volume Adjustment Factor | 1.0439 | 1.044 | 1.0262
Original Analytical Data
Pollutant | Sample ID | Sample Speciation (pg/m3) |  | Detection Limits (pg/m3)
|  | Run 1 | Run 2 | Run 1 | Run 2
2378-TCDD | Sample 1 | 24.96 | 22.29 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 8.46 | 8.46 | ND | ND
12378-PECDD | Sample 1 | 6.92 | 2.15 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 2.73 | 2.73 | ND | ND
123478-HXCDD | Sample 1 | 4.54 | 1.26 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 9.47 | 9.47 | ND | ND
123678-HXCDD | Sample 1 | 25.21 | 7.76 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 14.79 | 14.79 | ND | ND
123789-HXCDD | Sample 1 | 11.38 | 3.81 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 11.51 | 11.51 | ND | ND
1,2,3,4,6,7,8-HpCDD | Sample 1 | 131.49 | 39.24 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 90.61 | 90.61 | ND | ND
OCDD | Sample 1 | 344.67 | 138.71 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 789.83 | 789.83 | ND | ND
2,3,7,8-TCDF | Sample 1 | 31.172 | 9.946 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 9.404 | 9.404 | ND | ND
12378-PECDF | Sample 1 | 14.3 | 4.4 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 4.62 | 4.62 | ND | ND
23478-PECDF | Sample 1 | 14.94 | 3.89 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 4.3 | 4.3 | ND | ND
1,2,3,4,7,8-HxCDF | Sample 1 | 22.17 | 7.16 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 11.84 | 11.84 | ND | ND
1,2,3,6,7,8-HxCDF | Sample 1 | 9.21 | 3.31 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 6.77 | 6.77 | ND | ND
123789-HXCDF | Sample 1 | ND | ND | 13.84 | 3.53
| Sample 2 | ND | ND | ND | ND
| Background | 7.56 | 7.56 | ND | ND
2,3,4,6,7,8-HxCDF | Sample 1 | 5.04 | 1.71 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 10.15 | 10.15 | ND | ND
1,2,3,4,6,7,8-HpCDF | Sample 1 | 79.38 | 24.08 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 48.99 | 48.99 | ND | ND
1234789-HPCDF | Sample 1 | 2.32 | 2.32 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 39.55 | 39.55 | ND | ND
OCDF | Sample 1 | 17.87 | 13.8 | ND | ND
| Sample 2 | ND | ND | ND | ND
| Background | 240.64 | 240.64 | ND | ND

## Energetics
Energetics Data and Emission Factor Calculations
A068 - M196 5.56mm Tracer Cartridge, (Weapon was M16A1),
Test Run Number | 1 | 2 | Background
Number of Items Tested | 20 | 20 | 0
Net Explosive Weight for All Items (lb) | 0.0768 | 0.0768 | NA
Initial Plume Volume (m3) | 14.716 | 14.716 | 14.716
Dilution Correction Factor | 0.89 | 0.888 | NA
Is Sample 1 Active? | Yes | Yes | Yes
Is Sample 2 Active? | No | No | No
Energetic (Resin) Sample Volume (m3) Corrected to EPA Stationary Source Program Standard Conditions
Sample 1 | 0.58902 | 0.57371
Sample 2 | NA | NA
Background | 0.56198 | 0.56198
Pollutant | Sample ID | Sample Speciation (ug) |  | Detection Limits (ug) |  | Pollutant | Run Number | Sample Speciation with Detection Limits (SSDL) (ug)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC) (mg/m3)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3) | Test Concentration Relative Percent Difference
|  | Run 1 | Run 2 | Run 1 | Run 2 |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
Nitrobenzene | Sample 1 | ND | ND | 0.6 | 0.6 | Nitrobenzene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0010186 | Not Applicable | 0.0010186 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0010458 | Not Applicable
| Background | ND | ND | 0.6 | 0.6 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2-Nitrotoluene | Sample 1 | ND | ND | 0.6 | 0.6 | 2-Nitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0010186 | Not Applicable | 0.0010186 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0010458 | Not Applicable
| Background | ND | ND | 0.6 | 0.6 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
3-Nitrotoluene | Sample 1 | ND | ND | 0.6 | 0.6 | 3-Nitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0010186 | Not Applicable | 0.0010186 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0010458 | Not Applicable
| Background | ND | ND | 0.6 | 0.6 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
4-Nitrotoluene | Sample 1 | ND | ND | 0.6 | 0.6 | 4-Nitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0010186 | Not Applicable | 0.0010186 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0010458 | Not Applicable
| Background | ND | ND | 0.6 | 0.6 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Nitroglycerine | Sample 1 | ND | ND | 0.6 | 0.6 | Nitroglycerine | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0010186 | Not Applicable | 0.0010186 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0010458 | Not Applicable
| Background | ND | ND | 0.6 | 0.6 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,3-Dinitrobenzene | Sample 1 | ND | ND | 0.6 | 0.6 | 1,3-Dinitrobenzene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0010186 | Not Applicable | 0.0010186 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0010458 | Not Applicable
| Background | ND | ND | 0.6 | 0.6 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,6-Dinitrotoluene | Sample 1 | ND | ND | 0.6 | 0.6 | 2,6-Dinitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0010186 | Not Applicable | 0.0010186 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0010458 | Not Applicable
| Background | ND | ND | 0.6 | 0.6 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,4-Dinitrotoluene | Sample 1 | ND | ND | 0.6 | 0.6 | 2,4-Dinitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0010186 | Not Applicable | 0.0010186 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0010458 | Not Applicable
| Background | ND | ND | 0.6 | 0.6 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,3,5-Trinitrobenzene | Sample 1 | ND | ND | 0.6 | 0.6 | 1,3,5-Trinitrobenzene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0010186 | Not Applicable | 0.0010186 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0010458 | Not Applicable
| Background | ND | ND | 0.6 | 0.6 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,4,6-Trinitrotoluene | Sample 1 | ND | ND | 0.6 | 0.6 | 2,4,6-Trinitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0010186 | Not Applicable | 0.0010186 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0010458 | Not Applicable
| Background | ND | ND | 0.6 | 0.6 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
RDX | Sample 1 | ND | ND | 0.6 | 0.6 | RDX | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0010186 | Not Applicable | 0.0010186 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0010458 | Not Applicable
| Background | ND | ND | 0.6 | 0.6 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
4-Amino-2,6-Dinitrotoluene | Sample 1 | ND | ND | 0.6 | 0.6 | 4-Amino-2,6-Dinitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0010186 | Not Applicable | 0.0010186 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0010458 | Not Applicable
| Background | ND | ND | 0.6 | 0.6 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2-Amino-4,6-Dinitrotoluene | Sample 1 | ND | ND | 0.6 | 0.6 | 2-Amino-4,6-Dinitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0010186 | Not Applicable | 0.0010186 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0010458 | Not Applicable
| Background | ND | ND | 0.6 | 0.6 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Tetryl | Sample 1 | ND | ND | 0.6 | 0.6 | Tetryl | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0010186 | Not Applicable | 0.0010186 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0010458 | Not Applicable
| Background | ND | ND | 0.6 | 0.6 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
HMX | Sample 1 | ND | ND | 1.2 | 1.2 | HMX | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0020373 | Not Applicable | 0.0020373 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0020917 | Not Applicable
| Background | ND | ND | 1.2 | 1.2 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Energetic (Filter) Sample Volume (m3) Corrected to EPA Stationary Source Program Standard Conditions
| Run 1 | Run 2
Sample 1 | 0.58902 | 0.57371
Sample 2 | NA | NA
Background | 0.56198 | 0.56198
Pollutant | Sample ID | Sample Speciation (ug) |  | Detection Limits (ug) |  | Pollutant | Run Number | Sample Speciation with Detection Limits (SSDL) (ug)a |  | Background Corrected Concentration (BCC) (mg/m3)b |  | Dilution Corrected Concentration (DCC) (mg/m3)c |  | Quantity of Pollutant Emitted (QPE) (lb)d |  | Emission Factor (lb/item)e |  | Emission Factor (lb/lb NEW)f |  | Average Emission Factor |  | Detection Limit Based Concentration (mg/m3)g |  | Minimum Detection Limit Based Concentration (mg/m3) | Test Concentration Relative Percent Difference
|  | Run 1 | Run 2 | Run 1 | Run 2 |  |  | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | Sample 1 | Sample 2 | (lb/item) | (lb/lb NEW) | Train A | Train B
Nitrobenzene | Sample 1 | ND | ND | 0.4 | 0.4 | Nitrobenzene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00067909 | Not Applicable | 0.00067909 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00069722 | Not Applicable
| Background | ND | ND | 0.4 | 0.4 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2-Nitrotoluene | Sample 1 | ND | ND | 0.4 | 0.4 | 2-Nitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00067909 | Not Applicable | 0.00067909 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00069722 | Not Applicable
| Background | ND | ND | 0.4 | 0.4 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
3-Nitrotoluene | Sample 1 | ND | ND | 0.4 | 0.4 | 3-Nitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00067909 | Not Applicable | 0.00067909 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00069722 | Not Applicable
| Background | ND | ND | 0.4 | 0.4 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
4-Nitrotoluene | Sample 1 | ND | ND | 0.4 | 0.4 | 4-Nitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00067909 | Not Applicable | 0.00067909 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00069722 | Not Applicable
| Background | ND | ND | 0.4 | 0.4 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Nitroglycerine | Sample 1 | ND | ND | 1.2 | 0.4 | Nitroglycerine | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0020373 | Not Applicable | 0.00069722 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00069722 | Not Applicable
| Background | ND | ND | 0.4 | 0.4 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,3-Dinitrobenzene | Sample 1 | ND | ND | 0.4 | 0.4 | 1,3-Dinitrobenzene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00067909 | Not Applicable | 0.00067909 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00069722 | Not Applicable
| Background | ND | ND | 0.4 | 0.4 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,6-Dinitrotoluene | Sample 1 | ND | ND | 0.4 | 0.4 | 2,6-Dinitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00067909 | Not Applicable | 0.00067909 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00069722 | Not Applicable
| Background | ND | ND | 0.4 | 0.4 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,4-Dinitrotoluene | Sample 1 | ND | ND | 0.4 | 0.4 | 2,4-Dinitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00067909 | Not Applicable | 0.00067909 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00069722 | Not Applicable
| Background | ND | ND | 0.4 | 0.4 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
1,3,5-Trinitrobenzene | Sample 1 | ND | ND | 0.4 | 0.4 | 1,3,5-Trinitrobenzene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00067909 | Not Applicable | 0.00067909 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00069722 | Not Applicable
| Background | ND | ND | 0.4 | 0.4 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2,4,6-Trinitrotoluene | Sample 1 | ND | ND | 0.4 | 0.4 | 2,4,6-Trinitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00067909 | Not Applicable | 0.00067909 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00069722 | Not Applicable
| Background | ND | ND | 0.4 | 0.4 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
RDX | Sample 1 | ND | ND | 0.4 | 0.4 | RDX | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00067909 | Not Applicable | 0.00067909 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00069722 | Not Applicable
| Background | ND | ND | 0.4 | 0.4 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
4-Amino-2,6-Dinitrotoluene | Sample 1 | ND | ND | 0.4 | 0.4 | 4-Amino-2,6-Dinitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00067909 | Not Applicable | 0.00067909 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00069722 | Not Applicable
| Background | ND | ND | 0.4 | 0.4 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
2-Amino-4,6-Dinitrotoluene | Sample 1 | ND | ND | 0.4 | 0.4 | 2-Amino-4,6-Dinitrotoluene | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00067909 | Not Applicable | 0.00067909 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00069722 | Not Applicable
| Background | ND | ND | 0.4 | 0.4 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
Tetryl | Sample 1 | ND | ND | 0.4 | 0.4 | Tetryl | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.00067909 | Not Applicable | 0.00067909 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.00069722 | Not Applicable
| Background | ND | ND | 0.4 | 0.4 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
HMX | Sample 1 | ND | ND | 4 | 4 | HMX | 1 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | 0.0067909 | Not Applicable | 0.0067909 | ND
| Sample 2 | ND | ND | ND | ND |  | 2 | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |  |  | 0.0069722 | Not Applicable
| Background | ND | ND | 4 | 4 |  | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- | -- |  |  | -- | --
a If the corresponding sample speciation value does not = "ND" or all active sample speciation values = "ND," Sample Speciation with Detection Limits (SSDL) = Sample Speciation. Otherwise, (i.e., if the corresponding sample speciation value = "ND" and
at least one of the active sample speciation values does not = "ND," Sample Speciation with Detection Limits (SSDL) = 0.5 * Detection Limit.
b Background Corrected Concentration (BCC) = Test Mass/Test Sample Volume - Background Mass/ Background Sample Volume
c Dilution-Corrected Concentration (DCC) = BCC / Dilution Correction Factor
d Quantity of Pollutant Emitted (QPE) =BDCC * Initial Plume Volume
e Emission Factor = QPE / No. of Items
f Emission Factor = QPE / Net Explosive Weight
g If Average Emission Factor = "ND", Minimum Detection Limit Based Concentration = Test DL/Test Sample volume in mg/m3
[6 empty] | Pollutant | Total Emission Factorh |  | Overall Detection Limit Based Concentration (mg/m3)i
[7 empty] | (lb/item) | (lb/lb NEW)
[6 empty] | Nitrobenzene | ND | ND | 0.00067909
[6 empty] | 2-Nitrotoluene | ND | ND | 0.00067909
[6 empty] | 3-Nitrotoluene | ND | ND | 0.00067909
[6 empty] | 4-Nitrotoluene | ND | ND | 0.00067909
[6 empty] | Nitroglycerine | ND | ND | 0.00069722
[6 empty] | 1,3-Dinitrobenzene | ND | ND | 0.00067909
[6 empty] | 2,6-Dinitrotoluene | ND | ND | 0.00067909
[6 empty] | 2,4-Dinitrotoluene | ND | ND | 0.00067909
[6 empty] | 1,3,5-Trinitrobenzene | ND | ND | 0.00067909
[6 empty] | 2,4,6-Trinitrotoluene | ND | ND | 0.00067909
[6 empty] | RDX | ND | ND | 0.00067909
[6 empty] | 4-Amino-2,6-Dinitrotoluene | ND | ND | 0.00067909
[6 empty] | 2-Amino-4,6-Dinitrotoluene | ND | ND | 0.00067909
[6 empty] | Tetryl | ND | ND | 0.00067909
[6 empty] | HMX | ND | ND | 0.0067909
h Total Emission factor = Resin Emission Factor + Filter Emission Factor
i If Total Emission factor is "ND", Overall Detection Limit Based Concentration = Filter Minimum Detection Limit Based Concentration